## Cover Sheet
Analytical Cover Sheet
QAPP: | In Vitro TK Assessments of PFAS (D-CED-0031343)
Analysis Set ID: | 2021_PFAS_HepCl_G5_ALK | *Hepatic clearance assay of PFAS acids
Analytes: | DTXSID0059871 | PFAS
Dates Prepared: | 2021-06-02 | HepCl
Assay Preparer: | LDA, MJP
Dates Prepared: | 2021-06-02 | HepCl_CC
Dates of MS Run: | 2021-06-30
Assay Preparer: | ALK
Report Preparer: | ALK
Sample Type: | Standards, QCs, Unknown Analyte Samples | CC1-CC15
Internal Standard (conc) | 13C6-4-Nitrotoluene [lot SDFK-011] (25 pg/uL), 2-Perfluorooctyl-[1,1-2H2]-[1,2-13C2]-ethanol (8:2) [lot ] (25 pg/uL) | Fluorotelomers; Ref
MDL Calculation |  |  |  | References:
|  |  |  | 40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016)
eLOQ Definition | Lowest curve concentration at which precision checks can be quantitated against curve ±30%
QC Samples and Acceptance Criteria:
QC Type | Description | Acceptance Criteria | Type
QC Sample Type | Frequency | Acceptance Criteria | QC Metric | Corrective Action
(Analytical, TK Assay, or Both) |  |  | Assessed | if QC Fails
Calibration Curve Linearity (Analytical) | 1 each run | r2≥0.98; | Linearity | Rerun samples and curve
7-pt curve min.
Curve Check (Analytical) | 2-3 samples per run | 75-125 % | Accuracy | Flag in comments
Instrument Blanks (Analytical) | 1 every 6-10 injections | <½ LOD of instrument method | Specificity | Flag in comments, assess and resolve
Technical Replicate assessment | Every assay sample | 75-125% | Precision | Flag in comments
Reference Chemical(s) (Assay) | 1 per assay batch | RSD = ±20% of historical or published values | Reproducibility | Flag in comments
Analyte-IS Matching
Analyte | Name | Sample ID | IS | Reference Compound | Avg. MW
DTXSID0059871 | Pentafluoropropionamide | 273 | MFOET | Ametryn | 163.047
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | 900 | MFOET | Ametryn | 232.072
Sample ID Key
CC | Calibration Curve
QC | Quality Check
HT | Heat-treated Inactivated
Review of Data Timeline
Person | Action
LDA, MP | CC prepped 6/2/21
LDA, MP | Assay conducted 6/2/21
ALK | Samples prepped & run 6/30/21
MSC | Analytical data generated 7/15/21
ALK | QC Report completed 8/16/21
Notes
Concentrations listed are pre-crash concentrations in the assay
900 is still undergoing evaluation due to concerns over stability, as a 91% loss was seen in the HepCl assay in the Media controls by the T240 timepoint
Hepatic Clearance |  |  |  |  | Bkgd-Sub Clearance (µL/min/106 hepatocytes)
Assay Date | MS Date | Species | Compound | Conc. (nM) |  | Flags
2020-12-02 | 2021-06-30 | Human | 273 | 1000 | 0 | accuracy in CC, missing several samples
Stability in Williams' Media
Assay Date | MS Date | Compound | % of Time 0 at 240 min Avg | % of Time 0 at 240 min SD | Flags
2021-06-08 | 2021-06-09 | 273 | 47.77% | 11.21% | -
2021-06-08 | 2021-06-09 | 900 | 399.7% | 45.87% | -
Quantitative Limits
DTXSID | Name | Sample ID | MDL (nM) | eLOQ (nM)
DTXSID0059871 | Pentafluoropropionamide | 273 | 11.02 | 50

## Hep Clearance Calcs
Hepatocyte metabolic stabilty assay - Time course data and Clint calculations from Prism
Linear Regression Results
Analyte | Name | Sample ID | IS | Reference Compound | % Abiotic Loss* | % Abiotic Loss StDev | Concentration (nM) | Best-Fit Values | [8 empty] | No Cells BkgdCl |  |  |  | HT BkgdCl |  |  |  | Bkgd-Sub Clearance (µL/min/106 hepatocytes)
[5 empty] | (Mean +/- SD) | (Mean +/- SD) |  | Slope | r2 | Sy.x | F | DFn | DFd | P value | Half Life | Clearance | T-test pvalue | Slope | Half Life | Clearance | T-test pvalue | Slope | Half Life | Clearance
DTXSID0059871 | Pentafluoropropionamide | 273 | 4NT-13C6 | 4NT | 90.54% | 1.025% | 1000 | -0.00227 | 0.264 | 0.14 | 3.588 | 1 | 10 | 0.0875 | NS | NS | 0.0198 | -0.00383 | 180.9399 | 7.66 | 0.0003 | -0.003162 | 219.1651 | 6.324 | 0
| Ametryn | Ameryn |  |  | 83.56% | 28.43% |  | -0.01494 | 0.8039 | 0.6444 | 65.59 | 1 | 16 | <0.0001 | 46.3855 | 29.88 | 0.0059 | -0.002942 | 235.554 | 5.884 | <0.0001 | -0.002176 | 318.4743 | 4.352 | 23.996
Clint calcs:
k
0

## Hep Data for Prism
3/23/20 LC Runs of 3/4/20 & 3/6/20 Hepatocytes
Calcs
|  |  | 900 | [6 empty] | ln Transformed for Prism Entry- Clint Calcs on Hep Clear Data Tab
2020-12-02 | Human | 1uM | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | 900 | T=0 act | 1801.4588 | 103.6845 |  | 1737.442 | 6.781 | 1 uM | Y1 | Y2 | Y3
|  |  | 900 | T=0 act | 1601.4764 | 92.1744 |  |  |  | 0 | 7.4964 | 7.3787 | 7.5007
indicates <LOQ, so LOD/sqrt(2) used |  |  | 900 | T=0 act | 1809.3907 | 104.1411 |  |  |  | 15 | 6.7023 | 6.7544 | 6.6802
|  |  | 900 | T=15 act | 814.3003 | 98.9602 |  | 822.8566 | 3.833 | 30 | 6.5298 | 6.6649 | 6.4433
|  |  | 900 | T=15 act | 857.7919 | 104.2456 |  |  |  | 60 | 6.5844 | 6.6612 | 6.507
|  |  | 900 | T=15 act | 796.4775 | 96.7942 |  |  |  | 120 | 5.7745 | 5.8341 | 5.8324
|  |  | 900 | T=30 act | 685.2868 | 97.9828 |  | 699.3947 | 11.2807 | 240 | 4.6391 | 4.7057 | 4.8807
|  |  | 900 | T=30 act | 784.3937 | 112.1532
|  |  | 900 | T=30 act | 628.5036 | 89.8639
|  |  | 900 | T=60 act | 723.7133 | 99.8213 |  | 725.0089 | 7.7002
|  |  | 900 | T=60 act | 781.473 | 107.788
|  |  | 900 | T=60 act | 669.8406 | 92.3907
|  |  | 900 | T=120 act | 321.9754 | 96.1217 |  | 334.9664 | 3.3598
|  |  | 900 | T=120 act | 341.7518 | 102.0257
|  |  | 900 | T=120 act | 341.172 | 101.8526
|  |  | 900 | T=240 act | 103.4477 | 89.7609 |  | 115.2481 | 12.7615
|  |  | 900 | T=240 act | 110.5714 | 95.942
|  |  | 900 | T=240 act | 131.7252 | 114.2971
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD
|  |  | 900 | T=0 Inac | 1581.109 | 97.3727 |  | 1623.7696 | 10.0682 | Heat Inac
|  |  | 900 | T=0 Inac | 1485.8455 | 91.5059 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | 900 | T=0 Inac | 1804.3543 | 111.1213 |  |  |  | 0 | 7.3659 | 7.3037 | 7.498
|  |  | 900 | T=0 Media | 1774.2358 | 98.023 |  | 1810.0197 | 4.4675 | 240 | 5.0824 | 5.2084 | 5.123
|  |  | 900 | T=0 Media | 1902.6001 | 105.1149
|  |  | 900 | T=0 Media | 1753.2232 | 96.8621 |  |  |  | Media
|  |  | 900 | T=240 Inac | 161.1531 | 94.4642 |  | 170.5971 | 6.4975 | 1 uM | Y1 | Y2 | Y3
|  |  | 900 | T=240 Inac | 182.8009 | 107.1535 |  |  |  | 0 | 7.4811 | 7.551 | 7.4692
|  |  | 900 | T=240 Inac | 167.8374 | 98.3823 |  |  |  | 240 | 5.2415 | 5.1202 | 5.0574
|  |  | 900 | T=240 Media | 188.95 | 110.3928 | 89.35% | 171.1616 | 9.4796
|  |  | 900 | T=240 Media | 167.3607 | 97.7793 | 91.2% |  |  | % Abiotic Loss | 90.54% | % Abiotic Loss StDev | 1.025%
|  |  | 900 | T=240 Media | 157.1741 | 91.8279 | 91.04%
273
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | 273 | T=0 act | 1200.8952 | 94.9242 |  | 1265.109 | 6.112 | 1 uM | Y1 | Y2 | Y3
|  |  | 273 | T=0 act | 1350.9406 | 106.7845 |  |  |  | 0 | 7.0908 | 7.2086 | 7.1257
|  |  | 273 | T=0 act | 1243.4913 | 98.2912 |  |  |  | 15 | #NUM! | 6.8575 | 6.8595
|  |  | 273 | T=15 act |  | 0 |  | 951.9493 | 0.14367 | 30 | 6.8715 | 6.7808 | 6.8437
|  |  | 273 | T=15 act | 950.9822 | 99.8984 |  |  |  | 60 | 7.0037 | 6.9717 | 7.0853
|  |  | 273 | T=15 act | 952.9164 | 100.1016 |  |  |  | 120 | 6.6965 | #NUM! | #NUM!
|  |  | 273 | T=30 act | 964.3968 | 103.9536 |  | 927.7184 | 4.6068 | 240 | 6.3242 | 6.4199 | 6.4477
|  |  | 273 | T=30 act | 880.7866 | 94.9412
|  |  | 273 | T=30 act | 937.972 | 101.1052
|  |  | 273 | T=60 act | 1100.6659 | 98.2469 |  | 1120.3061 | 5.9201
|  |  | 273 | T=60 act | 1066.0205 | 95.1544
|  |  | 273 | T=60 act | 1194.2318 | 106.5987
|  |  | 273 | T=120 act | 809.5769 | 100 |  | 809.5769 | #DIV/0!
|  |  | 273 | T=120 act |  | 0
|  |  | 273 | T=120 act |  | 0
|  |  | 273 | T=240 act | 557.9043 | 92.8246 |  | 601.0308 | 6.3788
|  |  | 273 | T=240 act | 613.9375 | 102.1474
|  |  | 273 | T=240 act | 631.2507 | 105.028
|  |  | Compound | Type | Conc. (nM) | Precision | %Abiotic Loss | Avg Conc. (nM) | RSD
|  |  | 273 | T=0 Inac | 1278.2109 | 103.1118 |  | 1239.6358 | 10.1029 | Heat Inac
|  |  | 273 | T=0 Inac | 1099.6471 | 88.7073 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | 273 | T=0 Inac | 1341.0495 | 108.1809 |  |  |  | 0 | 7.1532 | 7.0027 | 7.2012
|  |  | 273 | T=0 Media |  | 0 |  | 1622.9047 | 0.92173 | 240 | 6.3127 | 6.4133 | 6.3544
|  |  | 273 | T=0 Media | 1612.3272 | 99.3482
|  |  | 273 | T=0 Media | 1633.4821 | 100.6518 |  |  |  | Media
|  |  | 273 | T=240 Inac | 551.557 | 95.287 |  | 578.8376 | 5.0736 | 1 uM | Y1 | Y2 | Y3
|  |  | 273 | T=240 Inac | 609.9219 | 105.3701 |  |  |  | 0 | #NUM! | 7.3854 | 7.3985
|  |  | 273 | T=240 Inac | 575.034 | 99.3429 |  |  |  | 240 | 6.3414 | #NUM! | 6.6038
|  |  | 273 | T=240 Media | 567.6166 | 86.9556 | #DIV/0! | 652.7657 | 18.4475
|  |  | 273 | T=240 Media |  | 0 | 100% |  |  | % Abiotic Loss* | 59.78% | % Abiotic Loss StDev | #DIV/0!
|  |  | 273 | T=240 Media | 737.9149 | 113.0444 | 54.83%
*% Abiotic Loss >50%, so use stability data to background subtract
Ametryn
|  |  | Compound | Type | Conc. (nM) | Precision |  | Avg Conc. (nM) | RSD | Human
|  |  | Ametryn | T=0 act | 1551.875 | 103.7064 |  | 1496.4126 | 4.509 | 1 uM | Y1 | Y2 | Y3
|  |  | Ametryn | T=0 act | 1421.2946 | 94.9801 |  |  |  | 0 | 7.3472 | 7.2593 | 7.3239
|  |  | Ametryn | T=0 act | 1516.0682 | 101.3135 |  |  |  | 15 | 6.8041 | 6.8773 | 6.8721
|  |  | Ametryn | T=15 act | 901.5677 | 95.3508 |  | 945.5273 | 4.035 | 30 | 6.8223 | 6.7356 | 6.8342
|  |  | Ametryn | T=15 act | 970.0039 | 102.5887 |  |  |  | 60 | 6.8454 | 6.8302 | 6.9061
|  |  | Ametryn | T=15 act | 965.0104 | 102.0606 |  |  |  | 120 | 7.1091 | 7.153 | 6.1677
|  |  | Ametryn | T=30 act | 918.0819 | 102.427 |  | 896.3279 | 5.3031 | 240 | 3.7435 | 3.0103 | 3.2247
|  |  | Ametryn | T=30 act | 841.8102 | 93.9177
|  |  | Ametryn | T=30 act | 929.0916 | 103.6553
|  |  | Ametryn | T=60 act | 939.5136 | 98.4402 |  | 954.4003 | 4.0543
|  |  | Ametryn | T=60 act | 925.3606 | 96.9573
|  |  | Ametryn | T=60 act | 998.3266 | 104.6025
|  |  | Ametryn | T=120 act | 1223.0592 | 123.205 |  | 992.7027 | 45.0653
|  |  | Ametryn | T=120 act | 1277.9451 | 128.7339
|  |  | Ametryn | T=120 act | 477.1036 | 48.0611
|  |  | Ametryn | T=240 act | 42.2453 | 144.5348 |  | 29.2284 | 39.4516
|  |  | Ametryn | T=240 act | 20.2935 | 69.4308
|  |  | Ametryn | T=240 act | 25.1465 | 86.0343
|  |  | Compound | Type | Conc. (nM) | Precision | %Abiotic Loss | Avg Conc. (nM) | RSD
|  |  | Ametryn | T=0 Inac | 1435.7885 | 99.3676 |  | 1444.9269 | 0.88086 | Heat Inac
|  |  | Ametryn | T=0 Inac | 1439.528 | 99.6264 |  |  |  | 1 uM | Y1 | Y2 | Y3
|  |  | Ametryn | T=0 Inac | 1459.4643 | 101.0061 |  |  |  | 0 | 7.2695 | 7.2721 | 7.2858
|  |  | Ametryn | T=0 Media | 1739.6641 | 104.6304 |  | 1662.6756 | 4.7079 | 240 | 6.7568 | 6.7506 | 6.7533
|  |  | Ametryn | T=0 Media | 1583.1721 | 95.2183
|  |  | Ametryn | T=0 Media | 1665.1907 | 100.1513 |  |  |  | Media
|  |  | Ametryn | T=240 Inac | 859.9235 | 100.3267 |  | 857.1234 | 0.314 | 1 uM | Y1 | Y2 | Y3
|  |  | Ametryn | T=240 Inac | 854.5558 | 99.7004 |  |  |  | 0 | 7.4614 | 7.3672 | 7.4177
|  |  | Ametryn | T=240 Inac | 856.8909 | 99.9729 |  |  |  | 240 | #NUM! | #NUM! | 6.7093
|  |  | Ametryn | T=240 Media | 0 | 0 | 100% | 273.3362 | 173.2051
|  |  | Ametryn | T=240 Media | 0 | 0 | 100% |  |  | % Abiotic Loss | 83.56% | % Abiotic Loss StDev | 28.43%
|  |  | Ametryn | T=240 Media | 820.0087 | 300 | 50.76%

## HC Data
Sample | [7 empty] | 900 Method | 900 Results |  |  |  | MFOET (ISTD) Results |  | 273 Method | 273 Results |  |  |  | MFOET (ISTD) Results
| Name | Comment | Info. | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | Final Conc. | RT | Area | Accuracy | RT | Resp. | Exp. Conc. | Final Conc. | RT | Area | Accuracy | RT | Resp.
| Hep5 CC15 |  |  | G5063021007.D | Cal | 15 | 2021-06-30 11:15 | 5000 | 4800.7736 | 8.2313 | 1103226.1436 | 96.0155 | 6.4659 | 17855.862 | 5000 | 4887.9131 | 9.2002 | 829230.9863 | 97.7583 | 6.4659 | 17855.862
| Hep5 CC14 |  |  | G5063021008.D | Cal | 14 | 2021-06-30 11:35 | 3500 | 3676.5908 | 8.2251 | 876443.7319 | 105.0455 | 6.459 | 18522.7832 | 3500 | 3625.6307 | 9.2001 | 650355.8056 | 103.5894 | 6.459 | 18522.7832
| Hep5 CC13 |  |  | G5063021009.D | Cal | 13 | 2021-06-30 11:55 | 2500 | 2601.0152 | 8.2252 | 645192.1724 | 104.0406 | 6.4582 | 19274.0819 | 2500 | 2531.7404 | 9.2002 | 480525.3448 | 101.2696 | 6.4582 | 19274.0819
| Hep5 CC12 |  |  | G5063021010.D | Cal | 12 | 2021-06-30 12:15 | 1500 | 1569.4184 | 8.2251 | 396727.1517 | 104.6279 | 6.459 | 19641.7735 | 1500 | 1538.8852 | 9.2001 | 302489.8966 | 102.5923 | 6.459 | 19641.7735
| Hep5 CC11 |  |  | G5063021011.D | Cal | 11 | 2021-06-30 12:35 | 800 | 845.1601 | 8.2191 | 220597.5246 | 105.645 | 6.4582 | 20280.989 | 800 | 806.6815 | 9.2002 | 166184.9181 | 100.8352 | 6.4582 | 20280.989
| Hep5 CC10 |  |  | G5063021012.D | Cal | 10 | 2021-06-30 12:55 | 500 | 474.9778 | 8.219 | 113380.006 | 94.9956 | 6.459 | 18547.7246 | 500 | 429.9992 | 9.2001 | 82132.8932 | 85.9998 | 6.459 | 18547.7246
| Hep5 CC9 |  |  | G5063021013.D | Cal | 9 | 2021-06-30 13:14 | 350 | 317.3579 | 8.2191 | 82730.7967 | 90.6737 | 6.4582 | 20255.6138 | 350 | 301.4939 | 9.2002 | 63487.6304 | 86.1411 | 6.4582 | 20255.6138
| Hep5 CC8 |  |  | G5063021014.D | Cal | 8 | 2021-06-30 13:34 | 200 | 228.0756 | 8.219 | 55121.3513 | 114.0378 | 6.4513 | 18778.8313 | 200 | 209.6303 | 9.2001 | 41436.2157 | 104.8151 | 6.4513 | 18778.8313
| Hep5 CC7 |  |  | G5063021015.D | Cal | 7 | 2021-06-30 13:54 | 125 | 154.8017 | 8.2191 | 35993.3899 | 123.8413 | 6.4582 | 18066.5267 | 125 | 146.5011 | 9.2002 | 28320.8237 | 117.2009 | 6.4582 | 18066.5267
| Hep5 CC6 |  |  | G5063021016.D | Cal | 6 | 2021-06-30 14:14 | 80 | 77.9635 | 8.219 | 20016.3147 | 97.4544 | 6.4513 | 19948.9497 | 80 | 72.1548 | 9.2001 | 16231.7331 | 90.1935 | 6.4513 | 19948.9497
| Hep5 CC5 |  |  | G5063021017.D | Cal | 5 | 2021-06-30 14:34 | 50 | 45.1895 | 8.2191 | 11610.6163 | 90.379 | 6.4582 | 19963.8984 | 50 | 40.2015 | 9.2062 | 9765.5022 | 80.403 | 6.4582 | 19963.8984
| Hep5 CC4 |  |  | G5063021018.D | Cal | 4 | 2021-06-30 14:54 | 30 | 33.5656 | 8.219 | 8112.9347 | 111.8854 | 6.459 | 18780.6508 | 30 | 39.441 | 9.2001 | 9041.5894 | 131.4699 | 6.459 | 18780.6508
| Hep5 CC3 |  |  | G5063021019.D | Cal | 3 | 2021-06-30 15:14 | 20 | 26.7017 | 8.2191 | 6150.4708 | 133.5084 | 6.4582 | 17897.6957 | 20 | 28.1912 | 9.2062 | 6570.4542 | 140.9558 | 6.4582 | 17897.6957
| Hep5 CC2 |  |  | G5063021020.D | Cal | 2 | 2021-06-30 15:34 | 12 | 15.3322 | 8.219 | 3894.3884 | 127.7686 | 6.459 | 19736.1066 | 12 | 10.5251 | 9.2001 | 3700.8921 | 87.7089 | 6.459 | 19736.1066
| Hep5 CC1 |  |  | G5063021021.D | Cal | 1 | 2021-06-30 15:54 | 7 | 5.3596 | 8.2191 | 1356.6992 | 76.5652 | 6.4505 | 19668.9714 | 7 | 4.8384 | 9.2062 | 2550.8712 | 69.1205 | 6.4505 | 19668.9714
| Spike |  |  | G5063021001.D | MatrixBlank |  | 2021-06-30 09:15 |  | 0 | 7.5434 | 0 |  | 6.4435 | 21191.3587 |  | 0 | 9.0599 | 0 |  | 6.4435 | 21191.3587
| Spike |  |  | G5063021002.D | MatrixBlank |  | 2021-06-30 09:35 |  | 0 | 8.0365 | 0 |  | 6.4513 | 20918.9062 |  | 0 | 9.0662 | 0 |  | 6.4513 | 20918.9062
| Spike |  |  | G5063021003.D | MatrixBlank |  | 2021-06-30 09:55 |  | 0 | 8.3286 | 0 |  | 6.4505 | 18540.8316 |  | 0 | 9.3219 | 0 |  | 6.4505 | 18540.8316
| Spike |  |  | G5063021022.D | MatrixBlank |  | 2021-06-30 16:14 |  | 0 | 8.3224 | 0 |  | 6.459 | 19611.4341 |  | 0 | 9.2001 | 0 |  | 6.459 | 19611.4341
| Hep5 CC15 |  |  | G5063021004.D | QC | 15 | 2021-06-30 10:15 | 5000 | 4510.8345 | 8.2312 | 1173797.78 | 90.2167 | 6.4667 | 20219.1965 | 5000 | 4375.8733 | 9.2061 | 847156.5418 | 87.5175 | 6.4667 | 20219.1965
| Hep5 CC15 |  |  | G5063021005.D | QC | 15 | 2021-06-30 10:35 | 5000 | 4608.5509 | 8.2313 | 1153814.0997 | 92.171 | 6.4659 | 19453.5538 | 5000 | 4565.5466 | 9.2062 | 847971.9054 | 91.3109 | 6.4659 | 19453.5538
| Hep5 CC15 |  |  | G5063021006.D | QC | 15 | 2021-06-30 10:55 | 5000 | 4614.2357 | 8.2312 | 1169972.6058 | 92.2847 | 6.459 | 19701.6871 | 5000 | 4634.0397 | 9.2001 | 870769.1496 | 92.6808 | 6.459 | 19701.6871
| Hep5 CC7 |  |  | G5063021032.D | QC | 7 | 2021-06-30 19:33 | 125 | 134.1369 | 8.219 | 36781.1073 | 107.3095 | 6.459 | 21306.1022 | 125 | 126.2014 | 9.2001 | 29016.5062 | 100.9611 | 6.459 | 21306.1022
| Hep5 QCCC3 |  |  | G5063021042.D | QC | 3 | 2021-06-30 22:53 | 20 | 19.8959 | 8.219 | 5260.9686 | 99.4795 | 6.4513 | 20546.1133 | 20 | 16.9586 | 9.2001 | 5196.7657 | 84.7929 | 6.4513 | 20546.1133
| Hep5 CC10 |  |  | G5063021043.D | QC | 10 | 2021-06-30 23:13 | 500 | 412.8982 | 8.2191 | 109640.2387 | 82.5796 | 6.4582 | 20632.6255 | 500 | 401.6293 | 9.2002 | 85479.8594 | 80.3259 | 6.4582 | 20632.6255
| Hep5 CC4 |  |  | G5063021053.D | QC | 4 | 2021-07-01 02:32 | 30 | 28.6408 | 8.2191 | 7995.8204 | 95.4693 | 6.4582 | 21692.2979 | 30 | 32.8127 | 9.2002 | 8982.343 | 109.3757 | 6.4582 | 21692.2979
| Hep5 QCCC8 |  |  | G5063021054.D | QC | 8 | 2021-07-01 02:52 | 200 | 178.1492 | 8.219 | 47434.8393 | 89.0746 | 6.459 | 20689.0709 | 200 | 175.4289 | 9.2061 | 38492.2857 | 87.7145 | 6.459 | 20689.0709
| Hep5 CC8 |  |  | G5063021064.D | QC | 8 | 2021-07-01 06:11 | 200 | 203.5483 | 8.219 | 59980.3918 | 101.7742 | 6.459 | 22896.5098 | 200 | 217.9355 | 9.2001 | 52444.7921 | 108.9677 | 6.459 | 22896.5098
| Hep5 QCCC11 |  |  | G5063021065.D | QC | 11 | 2021-07-01 06:31 | 800 | 691.9835 | 8.2191 | 190630.8996 | 86.4979 | 6.4582 | 21405.4851 | 800 | 708.9022 | 9.2002 | 154556.0228 | 88.6128 | 6.4582 | 21405.4851
| Hep5 CC5 |  |  | G5063021075.D | QC | 5 | 2021-07-01 09:51 | 50 | 44.0234 | 8.2252 | 12980.8304 | 88.0467 | 6.4659 | 22911.1382 | 50 | 45.8163 | 9.2062 | 12514.0795 | 91.6326 | 6.4659 | 22911.1382
| Hep5 CC9 |  |  | G5063021085.D | QC | 9 | 2021-07-01 13:10 | 350 | 306.2524 | 8.2252 | 87645.3437 | 87.5007 | 6.4582 | 22237.0362 | 350 | 317.668 | 9.2002 | 73325.0276 | 90.7623 | 6.4582 | 22237.0362
| Hep5 CC6 |  |  | G5063021095.D | QC | 6 | 2021-07-01 16:29 | 80 | 69.1855 | 8.2191 | 21488.7509 | 86.4819 | 6.4505 | 24133.6794 | 80 | 83.7763 | 9.2002 | 22483.3122 | 104.7203 | 6.4505 | 24133.6794
| Hep5 CC9 |  |  | G5063021105.D | QC | 9 | 2021-07-01 19:49 | 350 | 296.5581 | 8.2191 | 91394.6336 | 84.7309 | 6.4582 | 23946.2964 | 350 | 288.1202 | 9.2002 | 71824.5523 | 82.3201 | 6.4582 | 23946.2964
| Hep5 CC4 |  |  | G5063021115.D | QC | 4 | 2021-07-01 23:08 | 30 | 27.3997 | 8.2191 | 9117.6563 | 91.3323 | 6.4505 | 25856.2196 | 30 | 44.7127 | 9.2002 | 13832.8091 | 149.0423 | 6.4505 | 25856.2196
| Hep5 CC12 |  |  | G5063021125.D | QC | 12 | 2021-07-02 02:27 | 1500 | 1315.8129 | 8.2252 | 476314.3014 | 87.7209 | 6.4582 | 28127.2314 | 1500 | 1389.9004 | 9.2002 | 392269.7916 | 92.66 | 6.4582 | 28127.2314
| 900 HepCl T0a |  |  | G5063021023.D | Sample |  | 2021-06-30 16:34 |  | 1801.4588 | 8.2191 | 464244.9188 |  | 6.4582 | 20023.9761 |  | 0 | 9.2549 | 0 |  | 6.4582 | 20023.9761
| 900 HepCl T0b |  |  | G5063021024.D | Sample |  | 2021-06-30 16:54 |  | 1601.4764 | 8.219 | 407235.6562 |  | 6.459 | 19758.4446 |  | 0 | 9.2548 | 0 |  | 6.459 | 19758.4446
| 900 HepCl T0c |  |  | G5063021025.D | Sample |  | 2021-06-30 17:14 |  | 1809.3907 | 8.2191 | 461222.393 |  | 6.4582 | 19806.3993 |  | 0 | 9.261 | 0 |  | 6.4582 | 19806.3993
| 900 HepCl T0a HITC |  |  | G5063021026.D | Sample |  | 2021-06-30 17:34 |  | 1581.109 | 8.219 | 445029.7448 |  | 6.459 | 21870.3001 |  | 0 | 9.0723 | 0 |  | 6.459 | 21870.3001
| 900 HepCl T0b HITC |  |  | G5063021027.D | Sample |  | 2021-06-30 17:54 |  | 1485.8455 | 8.2191 | 403191.1744 |  | 6.4582 | 21084.5753 |  | 0 | 9.261 | 0 |  | 6.4582 | 21084.5753
| 900 HepCl T0c HITC |  |  | G5063021028.D | Sample |  | 2021-06-30 18:14 |  | 1804.3543 | 8.219 | 463779.6238 |  | 6.459 | 19971.8066 |  | 0 | 9.0723 | 0 |  | 6.459 | 19971.8066
| 900 HepCl T0a Media |  |  | G5063021029.D | Sample |  | 2021-06-30 18:34 |  | 1774.2358 | 8.2191 | 482469.7471 |  | 6.4582 | 21129.3559 |  | 0 | 9.0724 | 0 |  | 6.4582 | 21129.3559
| 900 HepCl T0b Media |  |  | G5063021030.D | Sample |  | 2021-06-30 18:54 |  | 1902.6001 | 8.219 | 499137.6564 |  | 6.459 | 20384.5117 |  | 0 | 9.2548 | 0 |  | 6.459 | 20384.5117
| 900 HepCl T0c Media |  |  | G5063021031.D | Sample |  | 2021-06-30 19:13 |  | 1753.2232 | 8.2191 | 461686.6818 |  | 6.4582 | 20461.5085 |  | 0 | 9.261 | 0 |  | 6.4582 | 20461.5085
| 900 HepCl T15a |  |  | G5063021033.D | Sample |  | 2021-06-30 19:53 |  | 814.3003 | 8.2191 | 213077.6697 |  | 6.4582 | 20332.0337 |  | 0 | 9.261 | 0 |  | 6.4582 | 20332.0337
| 900 HepCl T15b |  |  | G5063021034.D | Sample |  | 2021-06-30 20:13 |  | 857.7919 | 8.219 | 222982.3282 |  | 6.459 | 20198.3547 |  | 0 | 9.2609 | 0 |  | 6.459 | 20198.3547
| 900 HepCl T15c |  |  | G5063021035.D | Sample |  | 2021-06-30 20:33 |  | 796.4775 | 8.2191 | 225618.6273 |  | 6.4582 | 22010.4506 |  | 0 | 9.261 | 0 |  | 6.4582 | 22010.4506
| 900 HepCl T30a |  |  | G5063021036.D | Sample |  | 2021-06-30 20:53 |  | 685.2868 | 8.219 | 193613.9682 |  | 6.4513 | 21952.8958 |  | 0 | 9.2548 | 0 |  | 6.4513 | 21952.8958
| 900 HepCl T30b |  |  | G5063021037.D | Sample |  | 2021-06-30 21:13 |  | 784.3937 | 8.2191 | 217164.5956 |  | 6.4582 | 21512.0823 |  | 0 | 9.2549 | 0 |  | 6.4582 | 21512.0823
| 900 HepCl T30c |  |  | G5063021038.D | Sample |  | 2021-06-30 21:33 |  | 628.5036 | 8.2251 | 171650.7837 |  | 6.459 | 21220.9847 |  | 0 | 9.2609 | 0 |  | 6.459 | 21220.9847
| 900 HepCl T60a |  |  | G5063021039.D | Sample |  | 2021-06-30 21:53 |  | 723.7133 | 8.2191 | 192266.4802 |  | 6.4582 | 20642.6074 |  | 0 | 9.2549 | 0 |  | 6.4582 | 20642.6074
| 900 HepCl T60b |  |  | G5063021040.D | Sample |  | 2021-06-30 22:13 |  | 781.473 | 8.219 | 208936.9385 |  | 6.459 | 20774.4128 |  | 0 | 9.2548 | 0 |  | 6.459 | 20774.4128
| 900 HepCl T60c |  |  | G5063021041.D | Sample |  | 2021-06-30 22:33 |  | 669.8406 | 8.2191 | 178120.4582 |  | 6.4582 | 20661.8814 |  | 0 | 9.2549 | 0 |  | 6.4582 | 20661.8814
| 900 HepCl T120a |  |  | G5063021044.D | Sample |  | 2021-06-30 23:33 |  | 321.9754 | 8.219 | 89206.6873 |  | 6.459 | 21527.9298 |  | 0 | 9.2001 | 0 |  | 6.459 | 21527.9298
| 900 HepCl T120b |  |  | G5063021045.D | Sample |  | 2021-06-30 23:53 |  | 341.7518 | 8.2191 | 92983.6913 |  | 6.4582 | 21140.9013 |  | 0 | 9.2062 | 0 |  | 6.4582 | 21140.9013
| 900 HepCl T120c |  |  | G5063021046.D | Sample |  | 2021-07-01 00:13 |  | 341.172 | 8.219 | 87319.1913 |  | 6.4513 | 19886.7517 |  | 0 | 9.2001 | 0 |  | 6.4513 | 19886.7517
| 900 HepCl T240a |  |  | G5063021047.D | Sample |  | 2021-07-01 00:32 |  | 103.4477 | 8.2191 | 28343.045 |  | 6.4505 | 21288.899 |  | 0 | 9.2002 | 0 |  | 6.4505 | 21288.899
| 900 HepCl T240b |  |  | G5063021048.D | Sample |  | 2021-07-01 00:52 |  | 110.5714 | 8.219 | 31027.5981 |  | 6.459 | 21803.8393 |  | 0 | 9.2001 | 0 |  | 6.459 | 21803.8393
| 900 HepCl T240c |  |  | G5063021049.D | Sample |  | 2021-07-01 01:12 |  | 131.7252 | 8.2191 | 36539.8431 |  | 6.4582 | 21553.8669 |  | 0 | 9.261 | 0 |  | 6.4582 | 21553.8669
| 900 HepCl T240a HITC |  |  | G5063021050.D | Sample |  | 2021-07-01 01:32 |  | 161.1531 | 8.2251 | 49653.8941 |  | 6.459 | 23940.9873 |  | 0 | 9.2609 | 0 |  | 6.459 | 23940.9873
| 900 HepCl T240b HITC |  |  | G5063021051.D | Sample |  | 2021-07-01 01:52 |  | 182.8009 | 8.2191 | 57858.362 |  | 6.4582 | 24593.2211 |  | 0 | 9.2549 | 0 |  | 6.4582 | 24593.2211
| 900 HepCl T240c HITC |  |  | G5063021052.D | Sample |  | 2021-07-01 02:12 |  | 167.8374 | 8.219 | 46667.7071 |  | 6.4513 | 21605.0516 |  | 0 | 9.2548 | 0 |  | 6.4513 | 21605.0516
| 900 HepCl T240a Media |  |  | G5063021055.D | Sample |  | 2021-07-01 03:12 |  | 188.95 | 8.2191 | 55486.4916 |  | 6.4582 | 22817.4862 |  | 0 | 9.2549 | 0 |  | 6.4582 | 22817.4862
| 900 HepCl T240b Media |  |  | G5063021056.D | Sample |  | 2021-07-01 03:32 |  | 167.3607 | 8.219 | 48962.3587 |  | 6.459 | 22731.9294 |  | 0 | 9.2609 | 0 |  | 6.459 | 22731.9294
| 900 HepCl T240c Media |  |  | G5063021057.D | Sample |  | 2021-07-01 03:52 |  | 157.1741 | 8.2191 | 45309.9824 |  | 6.4582 | 22399.6035 |  | 0 | 9.2002 | 0 |  | 6.4582 | 22399.6035
| 273 HepCl T0a |  |  | G5063021058.D | Sample |  | 2021-07-01 04:12 |  | 0 | 8.1581 | 0 |  | 6.459 | 22469.7639 |  | 1200.8952 | 9.2061 | 271718.3775 |  | 6.459 | 22469.7639
| 273 HepCl T0b |  |  | G5063021059.D | Sample |  | 2021-07-01 04:32 |  | 0 | 8.1522 | 0 |  | 6.4505 | 22575.2557 |  | 1350.9406 | 9.2002 | 306233.176 |  | 6.4505 | 22575.2557
| 273 HepCl T0c |  |  | G5063021060.D | Sample |  | 2021-07-01 04:52 |  | 0 | 8.1703 | 0 |  | 6.459 | 21961.0857 |  | 1243.4913 | 9.2001 | 274760.7201 |  | 6.459 | 21961.0857
| 273 HepCl T0a HITC |  |  | G5063021061.D | Sample |  | 2021-07-01 05:12 |  | 0 | 8.3225 | 0 |  | 6.4582 | 21953.9139 |  | 1278.2109 | 9.2002 | 282153.8786 |  | 6.4582 | 21953.9139
| 273 HepCl T0b HITC |  |  | G5063021062.D | Sample |  | 2021-07-01 05:32 |  | 0 | 8.3164 | 0 |  | 6.4513 | 23463.5475 |  | 1099.6471 | 9.2001 | 260340.1971 |  | 6.4513 | 23463.5475
| 273 HepCl T0c HITC |  |  | G5063021063.D | Sample |  | 2021-07-01 05:52 |  | 0 | 8.3286 | 0 |  | 6.4505 | 22322.6811 |  | 1341.0495 | 9.2002 | 300644.7127 |  | 6.4505 | 22322.6811
| 273 HepCl T0a Media |  |  | G5063021066.D | Sample |  | 2021-07-01 06:51 |  | 0 | 7.647 | 0 |  | 6.4974 | 0 |  | 0 | 9.3035 | 0 |  | 6.4974 | 0
| 273 HepCl T0b Media |  |  | G5063021067.D | Sample |  | 2021-07-01 07:11 |  | 0 | 8.0548 | 0 |  | 6.4582 | 21758.3482 |  | 1612.3272 | 9.2002 | 350631.9539 |  | 6.4582 | 21758.3482
| 273 HepCl T0c Media |  |  | G5063021068.D | Sample |  | 2021-07-01 07:31 |  | 0 | 8.219 | 0 |  | 6.4513 | 21975.9859 |  | 1633.4821 | 9.2001 | 358655.7845 |  | 6.4513 | 21975.9859
| 273 HepCl T15a |  |  | G5063021069.D | Sample |  | 2021-07-01 07:51 |  | 0 | 8.1461 | 0 |  | 6.4659 | 0 |  | 0 | 9.1636 | 0 |  | 6.4659 | 0
| 273 HepCl T15b |  |  | G5063021070.D | Sample |  | 2021-07-01 08:11 |  | 0 | 8.0973 | 0 |  | 6.4667 | 19004.0635 |  | 950.9822 | 9.2061 | 182937.3698 |  | 6.4667 | 19004.0635
| 273 HepCl T15c |  |  | G5063021071.D | Sample |  | 2021-07-01 08:31 |  | 0 | 7.9757 | 0 |  | 6.4582 | 21935.183 |  | 952.9164 | 9.2062 | 211573.1713 |  | 6.4582 | 21935.183
| 273 HepCl T30a |  |  | G5063021072.D | Sample |  | 2021-07-01 08:51 |  | 0 | 8.0486 | 0 |  | 6.459 | 798.2423 |  | 964.3968 | 9.2061 | 7790.0929 |  | 6.459 | 798.2423
| 273 HepCl T30b |  |  | G5063021073.D | Sample |  | 2021-07-01 09:11 |  | 0 | 8.2252 | 0 |  | 6.4582 | 21976.5165 |  | 880.7866 | 9.2002 | 196256.9629 |  | 6.4582 | 21976.5165
| 273 HepCl T30c |  |  | G5063021074.D | Sample |  | 2021-07-01 09:31 |  | 0 | 8.2251 | 0 |  | 6.459 | 23704.9101 |  | 937.972 | 9.2061 | 225133.7094 |  | 6.459 | 23704.9101
| 273 HepCl T60a |  |  | G5063021076.D | Sample |  | 2021-07-01 10:11 |  | 9.7446 | 8.2251 | 2797.8253 |  | 6.459 | 22309.2495 |  | 1100.6659 | 9.2061 | 247756.8229 |  | 6.459 | 22309.2495
| 273 HepCl T60b |  |  | G5063021077.D | Sample |  | 2021-07-01 10:31 |  | 10.907 | 8.2252 | 3107.1841 |  | 6.4582 | 22135.5307 |  | 1066.0205 | 9.2062 | 238260.67 |  | 6.4582 | 22135.5307
| 273 HepCl T60c |  |  | G5063021078.D | Sample |  | 2021-07-01 10:50 |  | 13.1136 | 8.219 | 3608.4855 |  | 6.459 | 21381.0436 |  | 1194.2318 | 9.2001 | 257151.723 |  | 6.459 | 21381.0436
| 273 HepCl T120a |  |  | G5063021079.D | Sample |  | 2021-07-01 11:10 |  | 17.744 | 8.2191 | 2589.7398 |  | 6.4505 | 11340.4763 |  | 809.5769 | 9.2062 | 93251.7793 |  | 6.4505 | 11340.4763
| 273 HepCl T120b |  |  | G5063021080.D | Sample |  | 2021-07-01 11:30 |  | 0 | 8.3833 | 0 |  | 6.6049 | 0 |  | 0 | 9.2974 | 0 |  | 6.6049 | 0
| 273 HepCl T120c |  |  | G5063021081.D | Sample |  | 2021-07-01 11:50 |  | 0 | 8.2191 | 0 |  | 6.4736 | 0 |  | 0 | 9.2671 | 0 |  | 6.4736 | 0
| 273 HepCl T240a |  |  | G5063021082.D | Sample |  | 2021-07-01 12:10 |  | 24.8396 | 8.2251 | 7360.7003 |  | 6.4667 | 23025.1651 |  | 557.9043 | 9.2001 | 131506.1922 |  | 6.4667 | 23025.1651
| 273 HepCl T240b |  |  | G5063021083.D | Sample |  | 2021-07-01 12:30 |  | 36.6972 | 8.2252 | 10007.9115 |  | 6.4659 | 21190.3455 |  | 613.9375 | 9.2062 | 132908.4284 |  | 6.4659 | 21190.3455
| 273 HepCl T240c |  |  | G5063021084.D | Sample |  | 2021-07-01 12:50 |  | 31.4437 | 8.219 | 8518.5763 |  | 6.459 | 21050.4546 |  | 631.2507 | 9.2001 | 135674.2172 |  | 6.459 | 21050.4546
| 273 HepCl T240a HITC |  |  | G5063021086.D | Sample |  | 2021-07-01 13:30 |  | 36.1551 | 8.2251 | 10562.7585 |  | 6.459 | 22700.5217 |  | 551.557 | 9.2001 | 128208.9247 |  | 6.459 | 22700.5217
| 273 HepCl T240B HITC |  |  | G5063021087.D | Sample |  | 2021-07-01 13:50 |  | 36.1573 | 8.2252 | 10863.1826 |  | 6.4582 | 23344.734 |  | 609.9219 | 9.2062 | 145483.635 |  | 6.4582 | 23344.734
| 273 HepCl T240c HITC |  |  | G5063021088.D | Sample |  | 2021-07-01 14:10 |  | 39.6957 | 8.219 | 12311.7109 |  | 6.459 | 24099.1932 |  | 575.034 | 9.2001 | 141773.5666 |  | 6.459 | 24099.1932
| 273 HepCl T240a Media |  |  | G5063021089.D | Sample |  | 2021-07-01 14:30 |  | 39.9245 | 8.2191 | 12839.2432 |  | 6.4505 | 24987.7589 |  | 567.6166 | 9.2002 | 145145.4697 |  | 6.4505 | 24987.7589
| 273 HepCl T240b Media |  |  | G5063021090.D | Sample |  | 2021-07-01 14:50 |  | 0 | 8.1947 | 0 |  | 6.5205 | 0 |  | 0 | 9.1636 | 0 |  | 6.5205 | 0
| 273 HepCl T240c Media |  |  | G5063021091.D | Sample |  | 2021-07-01 15:10 |  | 43.352 | 8.2191 | 12403.7131 |  | 6.4582 | 22231.5785 |  | 737.9149 | 9.2002 | 166950.1499 |  | 6.4582 | 22231.5785
| 273 HepCl T0a Media |  |  | G5063021092.D | Sample |  | 2021-07-01 15:30 |  | 0 | 8.4137 | 0 |  | 6.4513 | 23943.1297 |  | 0 | 9.1879 | 0 |  | 6.4513 | 23943.1297
| 273 HepCl T0b Media |  |  | G5063021093.D | Sample |  | 2021-07-01 15:50 |  | 0 | 8.2313 | 0 |  | 6.4582 | 23838.2734 |  | 0 | 9.1819 | 0 |  | 6.4582 | 23838.2734
| 273 HepCl T0c Media |  |  | G5063021094.D | Sample |  | 2021-07-01 16:09 |  | 0 | 8.2129 | 0 |  | 6.459 | 23680.9204 |  | 0 | 9.2548 | 0 |  | 6.459 | 23680.9204
| Ametryn HepCl T0a |  |  | G5063021096.D | Sample |  | 2021-07-01 16:49 |  | 0 | 8.2251 | 0 |  | 6.4513 | 23464.6278 |  | 0 | 9.2001 | 0 |  | 6.4513 | 23464.6278
| Ametryn HepCl T0b |  |  | G5063021097.D | Sample |  | 2021-07-01 17:09 |  | 0 | 8.4077 | 0 |  | 6.4505 | 25215.0514 |  | 0 | 9.1941 | 0 |  | 6.4505 | 25215.0514
| Ametryn HepCl T0c |  |  | G5063021098.D | Sample |  | 2021-07-01 17:29 |  | 0 | 8.0425 | 0 |  | 6.4436 | 9091.0614 |  | 0 | 9.2609 | 0 |  | 6.4436 | 9091.0614
| Ametryn HepCl T0a HITC |  |  | G5063021099.D | Sample |  | 2021-07-01 17:49 |  | 0 | 8.4017 | 0 |  | 6.4582 | 23316.6381 |  | 0 | 9.261 | 0 |  | 6.4582 | 23316.6381
| Ametryn HepCl T0b HITC |  |  | G5063021100.D | Sample |  | 2021-07-01 18:09 |  | 0 | 8.3468 | 0 |  | 6.4513 | 23932.549 |  | 0 | 9.2548 | 0 |  | 6.4513 | 23932.549
| Ametryn HepCl T0c HITC |  |  | G5063021101.D | Sample |  | 2021-07-01 18:29 |  | 4.1296 | 8.2191 | 1280.0163 |  | 6.4582 | 24084.4283 |  | 0 | 9.2002 | 0 |  | 6.4582 | 24084.4283
| Ametryn HepCl T0a Media |  |  | G5063021102.D | Sample |  | 2021-07-01 18:49 |  | 0 | 8.2312 | 0 |  | 6.459 | 24599.6201 |  | 0 | 9.1879 | 0 |  | 6.459 | 24599.6201
| Ametryn HepCl T0a Media |  |  | G5063021103.D | Sample |  | 2021-07-01 19:09 |  | 0 | 7.6288 | 0 |  | 6.4582 | 24993.0751 |  | 0 | 9.261 | 0 |  | 6.4582 | 24993.0751
| Ametryn HepCl T0a Media |  |  | G5063021104.D | Sample |  | 2021-07-01 19:29 |  | 0 | 8.073 | 0 |  | 6.459 | 24316.2009 |  | 0 | 9.1879 | 0 |  | 6.459 | 24316.2009
| Ametryn HepCl T15a |  |  | G5063021106.D | Sample |  | 2021-07-01 20:09 |  | 0 | 8.2129 | 0 |  | 6.459 | 24551.793 |  | 0 | 9.194 | 0 |  | 6.459 | 24551.793
| Ametryn HepCl T15b |  |  | G5063021107.D | Sample |  | 2021-07-01 20:29 |  | 0 | 7.9696 | 0 |  | 6.4582 | 24773.5019 |  | 0 | 9.2549 | 0 |  | 6.4582 | 24773.5019
| Ametryn HepCl T15c |  |  | G5063021108.D | Sample |  | 2021-07-01 20:48 |  | 0 | 8.2251 | 0 |  | 6.459 | 23989.2094 |  | 0 | 9.2001 | 0 |  | 6.459 | 23989.2094
| Ametryn HepCl T30a |  |  | G5063021109.D | Sample |  | 2021-07-01 21:08 |  | 0 | 8.213 | 0 |  | 6.4505 | 24501.6464 |  | 0 | 9.2002 | 0 |  | 6.4505 | 24501.6464
| Ametryn HepCl T30b |  |  | G5063021110.D | Sample |  | 2021-07-01 21:28 |  | 0 | 8.2251 | 0 |  | 6.4513 | 26725.2233 |  | 0 | 9.2609 | 0 |  | 6.4513 | 26725.2233
| Ametryn HepCl T30c |  |  | G5063021111.D | Sample |  | 2021-07-01 21:48 |  | 0 | 8.2191 | 0 |  | 6.4582 | 26166.2116 |  | 0 | 9.2002 | 0 |  | 6.4582 | 26166.2116
| Ametryn HepCl T60a |  |  | G5063021112.D | Sample |  | 2021-07-01 22:08 |  | 4.7053 | 8.219 | 1560.2751 |  | 6.4513 | 25765.5073 |  | 0 | 9.2001 | 0 |  | 6.4513 | 25765.5073
| Ametryn HepCl T60b |  |  | G5063021113.D | Sample |  | 2021-07-01 22:28 |  | 6.9043 | 8.2191 | 2139.7551 |  | 6.4582 | 24080.7235 |  | 0 | 9.2002 | 0 |  | 6.4582 | 24080.7235
| Ametryn HepCl T60c |  |  | G5063021114.D | Sample |  | 2021-07-01 22:48 |  | 7.7287 | 8.219 | 2374.4011 |  | 6.4513 | 23871.3638 |  | 0 | 9.1879 | 0 |  | 6.4513 | 23871.3638
| Ametryn HepCl T120a |  |  | G5063021116.D | Sample |  | 2021-07-01 23:28 |  | 7.7545 | 8.219 | 2570.0333 |  | 6.4513 | 25752.0959 |  | 0 | 9.2001 | 0 |  | 6.4513 | 25752.0959
| Ametryn HepCl T120b |  |  | G5063021117.D | Sample |  | 2021-07-01 23:48 |  | 11.066 | 8.2191 | 3473.293 |  | 6.4582 | 24388.0253 |  | 0 | 9.2002 | 0 |  | 6.4582 | 24388.0253
| Ametryn HepCl T120c |  |  | G5063021118.D | Sample |  | 2021-07-02 00:08 |  | 11.4589 | 8.219 | 3825.3054 |  | 6.4513 | 25938.8137 |  | 0 | 9.2001 | 0 |  | 6.4513 | 25938.8137
| Ametryn HepCl T240a |  |  | G5063021119.D | Sample |  | 2021-07-02 00:28 |  | 6.6422 | 8.2191 | 2310.1468 |  | 6.4505 | 27024.2903 |  | 0 | 9.2002 | 0 |  | 6.4505 | 27024.2903
| Ametryn HepCl T240b |  |  | G5063021120.D | Sample |  | 2021-07-02 00:48 |  | 14.6951 | 8.219 | 4583.8933 |  | 6.459 | 24237.6299 |  | 0 | 9.2001 | 0 |  | 6.459 | 24237.6299
| Ametryn HepCl T240c |  |  | G5063021121.D | Sample |  | 2021-07-02 01:08 |  | 13.3528 | 8.2191 | 4430.3731 |  | 6.4582 | 25780.7572 |  | 0 | 9.2549 | 0 |  | 6.4582 | 25780.7572
| Ametryn HepCl T240a HITC |  |  | G5063021122.D | Sample |  | 2021-07-02 01:28 |  | 14.3987 | 8.219 | 5301.0132 |  | 6.4513 | 28606.3089 |  | 0 | 9.2001 | 0 |  | 6.4513 | 28606.3089
| Ametryn HepCl T240B HITC |  |  | G5063021123.D | Sample |  | 2021-07-02 01:48 |  | 0 | 8.2373 | 0 |  | 6.4813 | 0 |  | 0 | 9.2793 | 0 |  | 6.4813 | 0
| Ametryn HepCl T240c HITC |  |  | G5063021124.D | Sample |  | 2021-07-02 02:08 |  | 16.8708 | 8.219 | 5315.9309 |  | 6.459 | 24483.4254 |  | 0 | 9.2001 | 0 |  | 6.459 | 24483.4254

## AmetrynData
Sample | [8 empty] | Ametryn Method | Ametryn Results | [6 empty] | Qualifier (212.0 -> 71.1) Results |  | Qualifier (212.0 -> 122.0) Results |  | 4NT13C6 (ISTD) Results |  | Qualifier (143.0 -> 84.1) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | Ratio | MI | RT | Resp. | Ratio | MI
| HC_G5 CC 15 | E:\Shark Tank\RawData\Ametryn071421 |  | 52 | 7010_1_071421008.D | Cal | 15 | 2021-07-14 17:56 | 5000 | 16.8416 | 1447476.1368 | FALSE | 4929.8011 |  | 4929.8011 | 98.596 | 71.7894 | FALSE | 80.1471 | FALSE | 11.2876 | 1297527.8534 | 59.4584 | FALSE
| HC_G5 CC 14 | E:\Shark Tank\RawData\Ametryn071421 |  | 53 | 7010_1_071421009.D | Cal | 14 | 2021-07-14 18:20 | 3500 | 16.8372 | 1132024.6891 | FALSE | 3654.6558 |  | 3654.6558 | 104.4187 | 71.4421 | FALSE | 78.9524 | FALSE | 11.2874 | 1303947.0104 | 59.4224 | FALSE
| HC_G5 CC 13 | E:\Shark Tank\RawData\Ametryn071421 |  | 54 | 7010_1_071421010.D | Cal | 13 | 2021-07-14 18:44 | 2500 | 16.8415 | 814539.5518 | FALSE | 2472.6079 |  | 2472.6079 | 98.9043 | 71.385 | FALSE | 79.5992 | FALSE | 11.2875 | 1328423.813 | 59.0127 | FALSE
| HC_G5 CC 12 | E:\Shark Tank\RawData\Ametryn071421 |  | 55 | 7010_1_071421011.D | Cal | 12 | 2021-07-14 19:08 | 1500 | 16.8459 | 503415.2535 | FALSE | 1452.0577 |  | 1452.0577 | 96.8038 | 71.7307 | FALSE | 80.7133 | FALSE | 11.2876 | 1349036.6852 | 59.5179 | FALSE
| HC_G5 CC 11 | E:\Shark Tank\RawData\Ametryn071421 |  | 56 | 7010_1_071421012.D | Cal | 11 | 2021-07-14 19:31 | 800 | 16.8372 | 275703.0632 | FALSE | 796.1229 |  | 796.1229 | 99.5154 | 71.1097 | FALSE | 79.8666 | FALSE | 11.2831 | 1317849.3586 | 58.7339 | FALSE
| HC_G5 CC 10 | E:\Shark Tank\RawData\Ametryn071421 |  | 57 | 7010_1_071421013.D | Cal | 10 | 2021-07-14 19:55 | 500 | 16.8415 | 150316.853 | FALSE | 431.3599 |  | 431.3599 | 86.272 | 71.4051 | FALSE | 80.5203 | FALSE | 11.2875 | 1310034.4584 | 59.0358 | FALSE
| HC_G5 CC 9 | E:\Shark Tank\RawData\Ametryn071421 |  | 58 | 7010_1_071421014.D | Cal | 9 | 2021-07-14 20:20 | 350 | 16.8416 | 106533.851 | FALSE | 294.7516 |  | 294.7516 | 84.2147 | 72.5936 | FALSE | 81.4256 | FALSE | 11.2876 | 1352638.3831 | 59.184 | FALSE
| HC_G5 CC 8 | E:\Shark Tank\RawData\Ametryn071421 |  | 59 | 7010_1_071421015.D | Cal | 8 | 2021-07-14 20:44 | 200 | 16.8372 | 71453.1754 | FALSE | 203.803 |  | 203.803 | 101.9015 | 72.5107 | FALSE | 83.1051 | FALSE | 11.2874 | 1308151.8683 | 59.438 | FALSE
| HC_G5 CC 7 | E:\Shark Tank\RawData\Ametryn071421 |  | 60 | 7010_1_071421016.D | Cal | 7 | 2021-07-14 21:08 | 125 | 16.8415 | 47698.7009 | FALSE | 131.8852 |  | 131.8852 | 105.5082 | 70.9719 | FALSE | 81.0021 | FALSE | 11.2831 | 1346263.8751 | 58.2261 | FALSE
| HC_G5 CC 6 | E:\Shark Tank\RawData\Ametryn071421 |  | 61 | 7010_1_071421017.D | Cal | 6 | 2021-07-14 21:32 | 80 | 16.8416 | 28700.8737 | FALSE | 80.5536 |  | 80.5536 | 100.692 | 67.0924 | FALSE | 79.2784 | FALSE | 11.2876 | 1324029.6737 | 59.1845 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421018.D | Cal | 5 | 2021-07-14 21:56 | 50 | 16.8458 | 16333.6161 | FALSE | 45.9352 |  | 45.9352 | 91.8704 | 69.9999 | FALSE | 75.9474 | FALSE | 11.2874 | 1319870.783 | 58.2286 | FALSE
| HC_G5 CC 4 | E:\Shark Tank\RawData\Ametryn071421 |  | 63 | 7010_1_071421019.D | Cal | 4 | 2021-07-14 22:20 | 30 | 16.8458 | 10675.257 | FALSE | 29.5767 |  | 29.5767 | 98.589 | 65.3234 | FALSE | 86.0904 | FALSE | 11.2875 | 1339031.5749 | 58.7961 | FALSE
| HC_G5 CC 3 | E:\Shark Tank\RawData\Ametryn071421 |  | 64 | 7010_1_071421020.D | Cal | 3 | 2021-07-14 22:44 | 20 | 16.8373 | 7971.7468 | FALSE | 21.7899 |  | 21.7899 | 108.9497 | 67.8665 | FALSE | 78.4423 | FALSE | 11.2876 | 1356902.4864 | 58.6167 | FALSE
| HC_G5 CC 2 | E:\Shark Tank\RawData\Ametryn071421 |  | 65 | 7010_1_071421021.D | Cal | 2 | 2021-07-14 23:08 | 12 | 16.8415 | 4767.121 | FALSE | 13.0832 |  | 13.0832 | 109.0264 | 77.1353 | FALSE | 90.6814 | FALSE | 11.2831 | 1351047.4759 | 58.6343 | FALSE
| HC_G5 CC 1 | E:\Shark Tank\RawData\Ametryn071421 |  | 66 | 7010_1_071421022.D | Cal | 1 | 2021-07-14 23:32 | 7 | 16.8285 | 2398.0228 | FALSE | 6.6224 |  | 6.6224 | 94.6054 | 37.7605 | FALSE | 80.8701 | FALSE | 11.2831 | 1342378.2563 | 58.1003 | FALSE
| HC_G5 QCCC 3 | E:\Shark Tank\RawData\Ametryn071421 |  | 67 | 7010_1_071421023.D | Cal | 5 | 2021-07-14 23:55 | 50 | 16.8372 | 6543.7985 | FALSE | 17.1306 |  | 17.1306 | 34.2612 | 57.1286 | FALSE | 95.7227 | FALSE | 11.2874 | 1416584.6752 | 58.4806 | FALSE
| Method Spike | E:\Shark Tank\RawData\Ametryn071421 |  | 51 | 7010_1_071421001.D | MatrixBlank |  | 2021-07-14 15:09 |  | 16.8373 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2833 | 1204745.7807 | 59.7918 | FALSE
| Method Spike | E:\Shark Tank\RawData\Ametryn071421 |  | 51 | 7010_1_071421002.D | MatrixBlank |  | 2021-07-14 15:32 |  | 16.7077 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2874 | 1228409.96 | 59.2632 | FALSE
| Method Spike | E:\Shark Tank\RawData\Ametryn071421 |  | 51 | 7010_1_071421003.D | MatrixBlank |  | 2021-07-14 15:56 |  | 16.8329 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2875 | 1251806.8117 | 59.4666 | FALSE
| Method Spike | E:\Shark Tank\RawData\Ametryn071421 |  | 51 | 7010_1_071421004.D | MatrixBlank |  | 2021-07-14 16:20 |  | 16.876 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2874 | 1251698.649 | 58.8459 | FALSE
| Method Spike | E:\Shark Tank\RawData\Ametryn071421 |  | 1 | 7010_1_071421059.D | MatrixBlank |  | 2021-07-15 14:17 |  | 16.7854 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.322 | 374.3236 |  | FALSE
| Method Spike | E:\Shark Tank\RawData\Ametryn071421 |  | 1 | 7010_1_071421177.D | MatrixBlank |  | 2021-07-17 13:25 |  | 16.8156 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2745 | 1554859.589 | 59.9476 | FALSE
| HC_G5 CC 15 | E:\Shark Tank\RawData\Ametryn071421 |  | 52 | 7010_1_071421005.D | QC | 15 | 2021-07-14 16:44 | 5000 | 16.8372 | 1452857.6009 | FALSE | 4978.2559 |  | 4978.2559 | 99.5651 | 71.9004 | FALSE | 80.0713 | FALSE | 11.2875 | 1292118.1707 | 59.5952 | FALSE
| HC_G5 CC 15 | E:\Shark Tank\RawData\Ametryn071421 |  | 52 | 7010_1_071421006.D | QC | 15 | 2021-07-14 17:08 | 5000 | 16.8415 | 1464931.2346 | FALSE | 4995.5852 |  | 4995.5852 | 99.9117 | 71.5899 | FALSE | 79.9934 | FALSE | 11.2874 | 1299216.4988 | 59.5593 | FALSE
| HC_G5 CC 15 | E:\Shark Tank\RawData\Ametryn071421 |  | 52 | 7010_1_071421007.D | QC | 15 | 2021-07-14 17:32 | 5000 | 16.8415 | 1460810.52 | FALSE | 4974.5307 |  | 4974.5307 | 99.4906 | 71.1583 | FALSE | 79.76 | FALSE | 11.2875 | 1299974.8205 | 58.5855 | FALSE
| HC_G5 QCCC 11 | E:\Shark Tank\RawData\Ametryn071421 |  | 69 | 7010_1_071421033.D | QC | 11 | 2021-07-15 03:55 | 800 | 16.8458 | 238152.6754 | FALSE | 630.179 |  | 630.179 | 78.7724 | 70.5683 | FALSE | 80.1306 | FALSE | 11.2874 | 1430149.5922 | 58.1341 | FALSE
| HC_G5 CC 4 | E:\Shark Tank\RawData\Ametryn071421 |  | 63 | 7010_1_071421034.D | QC | 4 | 2021-07-15 04:19 | 30 | 16.8544 | 12860.7372 | FALSE | 34.602 |  | 34.602 | 115.3399 | 71.714 | FALSE | 79.6488 | FALSE | 11.2875 | 1379109.0049 | 58.4118 | FALSE
| HC_G5 CC 9 | E:\Shark Tank\RawData\Ametryn071421 |  | 58 | 7010_1_071421044.D | QC | 9 | 2021-07-15 08:18 | 350 | 16.8458 | 110679.1159 | FALSE | 293.5678 |  | 293.5678 | 83.8765 | 71.7035 | FALSE | 80.0632 | FALSE | 11.2875 | 1410881.3596 | 57.8701 | FALSE
| HC_G5 QCCC 8 | E:\Shark Tank\RawData\Ametryn071421 |  | 68 | 7010_1_071421045.D | QC | 8 | 2021-07-15 08:42 | 200 | 16.8501 | 60140.9393 | FALSE | 154.6367 |  | 154.6367 | 77.3183 | 71.9361 | FALSE | 82.6705 | FALSE | 11.2874 | 1448779.0359 | 57.5261 | FALSE
| HC_G5 CC 6 | E:\Shark Tank\RawData\Ametryn071421 |  | 61 | 7010_1_071421055.D | QC | 6 | 2021-07-15 12:41 | 80 | 16.8329 | 29457.1711 | FALSE | 75.4063 |  | 75.4063 | 94.2579 | 74.4025 | FALSE | 84.449 | FALSE | 11.2831 | 1451435.0126 | 57.5389 | FALSE
| HC_G5 CC 4 | E:\Shark Tank\RawData\Ametryn071421 |  | 63 | 7010_1_071421083.D | QC | 4 | 2021-07-15 23:53 | 30 | 16.8285 | 12790.5049 | FALSE | 33.0587 |  | 33.0587 | 110.1957 | 73.6747 | FALSE | 76.8818 | FALSE | 11.2788 | 1435533.068 | 59.0328 | FALSE
| HC_G5 CC 9 | E:\Shark Tank\RawData\Ametryn071421 |  | 58 | 7010_1_071421093.D | QC | 9 | 2021-07-16 03:52 | 350 | 16.8242 | 110341.576 | FALSE | 280.9628 |  | 280.9628 | 80.2751 | 71.0744 | FALSE | 79.9037 | FALSE | 11.2831 | 1469070.9574 | 60.3612 | FALSE
| HC_G5 CC 6 | E:\Shark Tank\RawData\Ametryn071421 |  | 61 | 7010_1_071421103.D | QC | 6 | 2021-07-16 07:52 | 80 | 16.82 | 30447.7688 | FALSE | 74.4055 |  | 74.4055 | 93.0068 | 68.231 | FALSE | 81.8542 | FALSE | 11.279 | 1520374.9655 | 60.8739 | FALSE
| G5-Ametryn Hep062221 T0a | E:\Shark Tank\RawData\Ametryn071421 |  | 70 | 7010_1_071421024.D | Sample |  | 2021-07-15 00:19 |  | 16.8458 | 547502.9618 | FALSE | 1551.875 |  | 1551.875 |  | 72.5378 | FALSE | 80.6845 | FALSE | 11.2875 | 1377535.2161 | 58.1504 | FALSE
| G5-Ametryn Hep062221 T0b | E:\Shark Tank\RawData\Ametryn071421 |  | 71 | 7010_1_071421025.D | Sample |  | 2021-07-15 00:43 |  | 16.8415 | 504854.2217 | FALSE | 1421.2946 |  | 1421.2946 |  | 70.3497 | FALSE | 79.7379 | FALSE | 11.2874 | 1380716.3354 | 58.593 | FALSE
| G5-Ametryn Hep062221 T0c | E:\Shark Tank\RawData\Ametryn071421 |  | 72 | 7010_1_071421026.D | Sample |  | 2021-07-15 01:07 |  | 16.8501 | 540722.9235 | FALSE | 1516.0682 |  | 1516.0682 |  | 70.5686 | FALSE | 80.4508 | FALSE | 11.2875 | 1390891.7473 | 58.1466 | FALSE
| G5-Ametryn Hep062221 Inactive T0a | E:\Shark Tank\RawData\Ametryn071421 |  | 73 | 7010_1_071421027.D | Sample |  | 2021-07-15 01:31 |  | 16.8416 | 527309.1094 | FALSE | 1435.7885 |  | 1435.7885 |  | 70.3049 | FALSE | 80.3777 | FALSE | 11.2876 | 1428280.2668 | 58.0177 | FALSE
| G5-Ametryn Hep062221 Inactive T0b | E:\Shark Tank\RawData\Ametryn071421 |  | 74 | 7010_1_071421028.D | Sample |  | 2021-07-15 01:55 |  | 16.8415 | 510982.0746 | FALSE | 1439.528 |  | 1439.528 |  | 71.6464 | FALSE | 80.5261 | FALSE | 11.2874 | 1380638.3707 | 58.2881 | FALSE
| G5-Ametryn Hep062221 Inactive T0c | E:\Shark Tank\RawData\Ametryn071421 |  | 75 | 7010_1_071421029.D | Sample |  | 2021-07-15 02:19 |  | 16.8415 | 514384.5203 | FALSE | 1459.4643 |  | 1459.4643 |  | 72.0679 | FALSE | 79.9568 | FALSE | 11.2875 | 1371785.4832 | 58.2129 | FALSE
| G5-Ametryn Hep062221 Media T0a | E:\Shark Tank\RawData\Ametryn071421 |  | 76 | 7010_1_071421030.D | Sample |  | 2021-07-15 02:43 |  | 16.8415 | 607281.0299 | FALSE | 1739.6641 |  | 1739.6641 |  | 69.9638 | FALSE | 79.6548 | FALSE | 11.2874 | 1371884.6679 | 58.2721 | FALSE
| G5-Ametryn Hep062221 Media T0b | E:\Shark Tank\RawData\Ametryn071421 |  | 77 | 7010_1_071421031.D | Sample |  | 2021-07-15 03:07 |  | 16.8458 | 577574.2165 | FALSE | 1583.1721 |  | 1583.1721 |  | 71.7092 | FALSE | 80.1288 | FALSE | 11.2875 | 1426006.1663 | 57.8877 | FALSE
| G5-Ametryn Hep062221 Media T0c | E:\Shark Tank\RawData\Ametryn071421 |  | 78 | 7010_1_071421032.D | Sample |  | 2021-07-15 03:31 |  | 16.8459 | 599082.2298 | FALSE | 1665.1907 |  | 1665.1907 |  | 70.3 | FALSE | 80.5836 | FALSE | 11.2876 | 1410246.6401 | 57.568 | FALSE
| G5-Ametryn Hep062221 T15a | E:\Shark Tank\RawData\Ametryn071421 |  | 79 | 7010_1_071421035.D | Sample |  | 2021-07-15 04:43 |  | 16.8502 | 337983.4307 | FALSE | 901.5677 |  | 901.5677 |  | 70.4425 | FALSE | 79.9612 | FALSE | 11.2876 | 1431668.5961 | 57.8436 | FALSE
| G5-Ametryn Hep062221 T15b | E:\Shark Tank\RawData\Ametryn071421 |  | 80 | 7010_1_071421036.D | Sample |  | 2021-07-15 05:07 |  | 16.8458 | 361463.7952 | FALSE | 970.0039 |  | 970.0039 |  | 70.4572 | FALSE | 79.4384 | FALSE | 11.2874 | 1426395.4767 | 57.9119 | FALSE
| G5-Ametryn Hep062221 T15c | E:\Shark Tank\RawData\Ametryn071421 |  | 81 | 7010_1_071421037.D | Sample |  | 2021-07-15 05:31 |  | 16.8458 | 355991.3632 | FALSE | 965.0104 |  | 965.0104 |  | 70.5194 | FALSE | 80.5791 | FALSE | 11.2875 | 1411831.2383 | 57.8989 | FALSE
| G5-Ametryn Hep062221 T30a | E:\Shark Tank\RawData\Ametryn071421 |  | 82 | 7010_1_071421038.D | Sample |  | 2021-07-15 05:55 |  | 16.8459 | 346576.8128 | FALSE | 918.0819 |  | 918.0819 |  | 70.7988 | FALSE | 80.4541 | FALSE | 11.2876 | 1442465.364 | 57.5782 | FALSE
| G5-Ametryn Hep062221 T30b | E:\Shark Tank\RawData\Ametryn071421 |  | 83 | 7010_1_071421039.D | Sample |  | 2021-07-15 06:19 |  | 16.8458 | 313684.4159 | FALSE | 841.8102 |  | 841.8102 |  | 70.6306 | FALSE | 80.2927 | FALSE | 11.2874 | 1420202.0068 | 58.3375 | FALSE
| G5-Ametryn Hep062221 T30c | E:\Shark Tank\RawData\Ametryn071421 |  | 84 | 7010_1_071421040.D | Sample |  | 2021-07-15 06:43 |  | 16.8501 | 340000.7099 | FALSE | 929.0916 |  | 929.0916 |  | 70.1613 | FALSE | 79.4359 | FALSE | 11.2875 | 1398845.959 | 58.0554 | FALSE
| G5-Ametryn Hep062221 T60a | E:\Shark Tank\RawData\Ametryn071421 |  | 85 | 7010_1_071421041.D | Sample |  | 2021-07-15 07:07 |  | 16.8458 | 350445.0242 | FALSE | 939.5136 |  | 939.5136 |  | 71.4602 | FALSE | 81.2426 | FALSE | 11.2874 | 1426324.0181 | 57.9673 | FALSE
| G5-Ametryn Hep062221 T60b | E:\Shark Tank\RawData\Ametryn071421 |  | 86 | 7010_1_071421042.D | Sample |  | 2021-07-15 07:30 |  | 16.8501 | 343994.2464 | FALSE | 925.3606 |  | 925.3606 |  | 72.8963 | FALSE | 81.0547 | FALSE | 11.2875 | 1420803.762 | 57.7486 | FALSE
| G5-Ametryn Hep062221 T60c | E:\Shark Tank\RawData\Ametryn071421 |  | 87 | 7010_1_071421043.D | Sample |  | 2021-07-15 07:54 |  | 16.8501 | 353146.74 | FALSE | 998.3266 |  | 998.3266 |  | 69.3509 | FALSE | 79.5645 | FALSE | 11.2874 | 1355336.1883 | 57.8549 | FALSE
| G5-Ametryn Hep062221 T120a | E:\Shark Tank\RawData\Ametryn071421 |  | 88 | 7010_1_071421046.D | Sample |  | 2021-07-15 09:06 |  | 16.8458 | 409751.9442 | FALSE | 1223.0592 |  | 1223.0592 |  | 72.292 | FALSE | 80.861 | FALSE | 11.2875 | 1293456.5952 | 58.1 | FALSE
| G5-Ametryn Hep062221 T120b | E:\Shark Tank\RawData\Ametryn071421 |  | 89 | 7010_1_071421047.D | Sample |  | 2021-07-15 09:30 |  | 16.8501 | 432912.7563 | FALSE | 1277.9451 |  | 1277.9451 |  | 71.576 | FALSE | 79.4778 | FALSE | 11.2874 | 1310326.9542 | 57.7717 | FALSE
| G5-Ametryn Hep062221 T120c | E:\Shark Tank\RawData\Ametryn071421 |  | 90 | 7010_1_071421048.D | Sample |  | 2021-07-15 09:54 |  | 16.8458 | 182707.2926 | FALSE | 477.1036 |  | 477.1036 |  | 71.0427 | FALSE | 78.7502 | FALSE | 11.2875 | 1441842.2549 | 57.9092 | FALSE
| G5-Ametryn Hep062221 T240a | E:\Shark Tank\RawData\Ametryn071421 |  | 91 | 7010_1_071421049.D | Sample |  | 2021-07-15 10:18 |  | 16.8502 | 16755.1958 | FALSE | 42.2453 |  | 42.2453 |  | 78.7087 | FALSE | 86.9938 | FALSE | 11.2876 | 1472019.8975 | 58.0955 | FALSE
| G5-Ametryn Hep062221 T240b | E:\Shark Tank\RawData\Ametryn071421 |  | 92 | 7010_1_071421050.D | Sample |  | 2021-07-15 10:42 |  | 16.8458 | 7926.9795 | FALSE | 20.2935 |  | 20.2935 |  | 64.2006 | FALSE | 77.2026 | FALSE | 11.2874 | 1448704.5787 | 57.5473 | FALSE
| G5-Ametryn Hep062221 T240c | E:\Shark Tank\RawData\Ametryn071421 |  | 93 | 7010_1_071421051.D | Sample |  | 2021-07-15 11:06 |  | 16.8588 | 9930.0925 | FALSE | 25.1465 |  | 25.1465 |  | 54.7873 | FALSE | 71.5113 | FALSE | 11.2875 | 1464788.2941 | 57.2922 | FALSE
| G5-Ametryn Hep062221 Inactive T240a | E:\Shark Tank\RawData\Ametryn071421 |  | 94 | 7010_1_071421052.D | Sample |  | 2021-07-15 11:30 |  | 16.8458 | 320533.1032 | FALSE | 859.9235 |  | 859.9235 |  | 70.7212 | FALSE | 81.7924 | FALSE | 11.2874 | 1421507.6497 | 58.1738 | FALSE
| G5-Ametryn Hep062221 Inactive T240b | E:\Shark Tank\RawData\Ametryn071421 |  | 95 | 7010_1_071421053.D | Sample |  | 2021-07-15 11:54 |  | 16.8415 | 324255.1247 | FALSE | 854.5558 |  | 854.5558 |  | 71.0053 | FALSE | 80.6014 | FALSE | 11.2831 | 1446785.1497 | 57.8535 | FALSE
| G5-Ametryn Hep062221 Inactive T240c | E:\Shark Tank\RawData\Ametryn071421 |  | 96 | 7010_1_071421054.D | Sample |  | 2021-07-15 12:18 |  | 16.8415 | 333970.3912 | FALSE | 856.8909 |  | 856.8909 |  | 72.4551 | FALSE | 81.1114 | FALSE | 11.2831 | 1486189.5174 | 58.1446 | FALSE
| G5-Ametryn Hep062221 Media T240a | E:\Shark Tank\RawData\Ametryn071421 |  | 97 | 7010_1_071421056.D | Sample |  | 2021-07-15 13:05 |  | 16.9019 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2874 | 1621893.8021 | 58.2267 | FALSE
| G5-Ametryn Hep062221 Media T240b | E:\Shark Tank\RawData\Ametryn071421 |  | 98 | 7010_1_071421057.D | Sample |  | 2021-07-15 13:29 |  | 16.9969 | 0 | TRUE | 0 |  | 0 |  |  | FALSE |  | FALSE | 11.2831 | 1607395.8666 | 57.7737 | FALSE
| G5-Ametryn Hep062221 Media T240c | E:\Shark Tank\RawData\Ametryn071421 |  | 99 | 7010_1_071421058.D | Sample |  | 2021-07-15 13:53 |  | 16.8328 | 311840.6349 | FALSE | 820.0087 |  | 820.0087 |  | 69.9823 | FALSE | 81.0114 | FALSE | 11.2831 | 1448328.8326 | 57.3185 | FALSE
| HC_G5 QCCC 8 | E:\Shark Tank\RawData\Ametryn071421 |  | 18 | 7010_1_071421149.D | Sample |  | 2021-07-17 02:15 |  | 16.8156 | 48586.2643 | FALSE | 119.4851 |  | 119.4851 |  | 68.6053 | FALSE | 79.5783 | FALSE | 11.2745 | 1513013.0649 | 59.8159 | FALSE

## 900_stab
Sample | [7 empty] | 900 Results |  |  |  | MFOET (ISTD) Results
| Name | Comment | Info. | Data File | Type | Level | Acq. Date-Time | Final Conc. | RT | Area | Accuracy | RT | Resp. |  |  |  |  | 900 HC Stability
| Spike |  |  | G5060921001.D | MatrixBlank |  | 2021-06-09 16:25 | 0 | 7.9755 | 2286.5676 |  | 6.4512 | 48954.0392 | Hep Clearance Time Course Sample Results |  |  |  |  | % of Time 0
| Spike |  |  | G5060921002.D | MatrixBlank |  | 2021-06-09 16:45 | 0 | 8.3286 | 282.4543 |  | 6.4513 | 48159.3935 | Sample ID | Peak Area | Average | Standard Deviation | Time | Replicate | Average | SD
| Spike |  |  | G5060921003.D | MatrixBlank |  | 2021-06-09 17:05 | 0 | 8.213 | 376.6974 |  | 6.4506 | 49705.5305 | T0 rep 1 | 55573.7181 | 58000.232 | 7845.5812 | 0 | 95.82% | 100% | 0.13527
| Spike |  |  | G5060921004.D | MatrixBlank |  | 2021-06-09 17:25 | 0 | 8.292 | 270.097 |  | 6.4513 | 46806.429 | T0 rep 2 | 66772.4021 |  |  |  | 115.1%
| Spike |  |  | G5060921005.D | MatrixBlank |  | 2021-06-09 17:45 | 0 | 8.1948 | 435.9725 |  | 6.4582 | 47636.5873 | T0 rep 3 | 51654.576 |  |  |  | 89.06%
| Hep5 CC15 |  |  | G5060921006.D | QC | 15 | 2021-06-09 18:04 | 4850.407 | 8.2312 | 1198856.1079 | 97.0081 | 6.459 | 48065.3218 | T15 rep 1 | 118456.3612 | 114807.1268 | 7196.6994 | 15 | 204.2% | 197.9% | 0.12408
| Hep5 CC15 |  |  | G5060921007.D | QC | 15 | 2021-06-09 18:24 | 4745.7331 | 8.2313 | 1195386.0344 | 94.9147 | 6.4582 | 48976.9109 | T15 rep 2 | 119448.1725 |  |  |  | 205.9%
| Hep5 CC15 |  |  | G5060921008.D | QC | 15 | 2021-06-09 18:44 | 4670.2635 | 8.2312 | 1196983.8761 | 93.4053 | 6.4582 | 49830.0319 | T15 rep 3 | 106516.8466 |  |  |  | 183.6%
| Hep5 CC15 |  |  | G5060921009.D | QC | 15 | 2021-06-09 19:04 | 4543.2973 | 8.2313 | 1161115.1427 | 90.8659 | 6.4582 | 49679.1447 | T30 rep 1 | 161559.6814 | 157459.7922 | 5663.3391 | 30 | 278.6% | 271.5% | 0.097643
| Hep5 CC15 |  |  | G5060921010.D | Cal | 15 | 2021-06-09 19:24 | 4885.1854 | 8.2312 | 1155520.1532 | 97.7037 | 6.4582 | 45999.9838 | T30 rep 2 | 159821.9437 |  |  |  | 275.6%
| Hep5 CC14 |  |  | G5060921011.D | Cal | 14 | 2021-06-09 19:44 | 5294.3404 | 8.2374 | 1279087.5552 | 151.2669 | 6.4659 | 47005.2191 | T30 rep 3 | 150997.7516 |  |  |  | 260.3%
| Hep5 CC13 |  |  | G5060921012.D | Cal | 13 | 2021-06-09 20:04 | 2725.6924 | 8.2312 | 689483.2934 | 109.0277 | 6.4667 | 48966.5011 | T60 rep 1 | 176339.1352 | 188944.6307 | 10916.912 | 60 | 304% | 325.8% | 0.18822 | Time (min) | % of Time 0
| Hep5 CC12 |  |  | G5060921013.D | Cal | 12 | 2021-06-09 20:24 | 0 | 8.0548 | 405.3572 | 0 | 6.4582 | 52771.4594 | T60 rep 2 | 195318.65 |  |  |  | 336.8% |  |  | 0 | 95.82% | 115.1% | 89.06%
| Hep5 CC11 |  |  | G5060921014.D | Cal | 11 | 2021-06-09 20:44 | 855.8258 | 8.2251 | 221018.0443 | 106.9782 | 6.4582 | 48876.3755 | T60 rep 3 | 195176.107 |  |  |  | 336.5% |  |  | 15 | 204.2% | 205.9% | 183.6%
| Hep5 CC10 |  |  | G5060921015.D | Cal | 10 | 2021-06-09 21:04 | 463.4456 | 8.2252 | 116469.3176 | 92.6891 | 6.4659 | 46288.9036 | T120 rep 1 | 203289.1371 | 213789.6856 | 9807.0481 | 120 | 350.5% | 368.6% | 0.16909 | 30 | 278.6% | 275.6% | 260.3%
| Hep5 CC9 |  |  | G5060921016.D | Cal | 9 | 2021-06-09 21:24 | 309.7659 | 8.2251 | 84398.3694 | 88.5045 | 6.459 | 48770.2368 | T120 rep 2 | 222711.7514 |  |  |  | 384% |  |  | 60 | 304% | 336.8% | 336.5%
| Hep5 CC8 |  |  | G5060921017.D | Cal | 8 | 2021-06-09 21:44 | 201.5837 | 8.2252 | 55352.3485 | 100.7919 | 6.4659 | 47008.1778 | T120 rep 3 | 215368.1682 |  |  |  | 371.3% |  |  | 120 | 350.5% | 384% | 371.3%
| Hep5 CC7 |  |  | G5060921018.D | Cal | 7 | 2021-06-09 22:04 | 137.7507 | 8.2251 | 40957.5359 | 110.2006 | 6.4582 | 48117.9084 | T240 rep 1 | 244311.9681 | 231810.3824 | 26605.9208 | 240 | 421.2% | 399.7% | 0.45872 | 240 | 421.2% | 347% | 430.8%
| Hep5 CC6 |  |  | G5060921019.D | Cal | 6 | 2021-06-09 22:24 | 55.1531 | 8.2252 | 20392.5373 | 68.9413 | 6.4582 | 47540.2486 | T240 rep 2 | 201256.1373 |  |  |  | 347%
| Hep5 CC5 |  |  | G5060921020.D | Cal | 5 | 2021-06-09 22:44 | 24.3384 | 8.2251 | 12305.6442 | 48.6767 | 6.4667 | 45336.5558 | T240 rep 3 | 249863.0416 |  |  |  | 430.8%
| Hep5 CC4 |  |  | G5060921021.D | Cal | 4 | 2021-06-09 23:04 | 10.1361 | 8.2252 | 9212.5935 | 33.7869 | 6.4659 | 46334.7654
| Hep5 CC3 |  |  | G5060921022.D | Cal | 3 | 2021-06-09 23:24 | 0 | 8.2251 | 6001.8177 | 0 | 6.4658 | 47432.6339
| Hep5 CC2 |  |  | G5060921023.D | Cal | 2 | 2021-06-09 23:44 | 21.1808 | 8.2252 | 747.2768 | 176.5064 | 6.4505 | 2927.2003
| Hep5 CC1 |  |  | G5060921024.D | Cal | 1 | 2021-06-10 00:03 | 0 | 8.2251 | 2507.6671 | 0 | 6.4513 | 29362.9715
| Spike |  |  | G5060921025.D | MatrixBlank |  | 2021-06-10 00:23 | 0 | 8.1278 | 1074.8494 |  | 6.4582 | 48509.0713
| Hep5 CC5 |  |  | G5060921045.D | QC | 5 | 2021-06-10 07:02 | 26.2893 | 8.2252 | 13755.3331 | 52.5787 | 6.4506 | 48881.418
| Plasma Stab 900 T0a |  |  | G5060921046.D | Sample |  | 2021-06-10 07:22 | 194.2668 | 8.219 | 55573.7181 |  | 6.4513 | 48744.5722
| Plasma Stab 900 T0b |  |  | G5060921047.D | Sample |  | 2021-06-10 07:42 | 224.1557 | 8.2191 | 66772.4021 |  | 6.4582 | 51645.7672
| Plasma Stab 900 T0c |  |  | G5060921048.D | Sample |  | 2021-06-10 08:02 | 178.4728 | 8.2251 | 51654.576 |  | 6.459 | 48760.078
| Plasma Stab 900 T15a |  |  | G5060921049.D | Sample |  | 2021-06-10 08:22 | 451.2244 | 8.2191 | 118456.3612 |  | 6.4505 | 48277.3307
| Plasma Stab 900 T15b |  |  | G5060921050.D | Sample |  | 2021-06-10 08:42 | 441.002 | 8.219 | 119448.1725 |  | 6.4505 | 49740.8992
| Plasma Stab 900 T15c |  |  | G5060921051.D | Sample |  | 2021-06-10 09:02 | 401.4844 | 8.2191 | 106516.8466 |  | 6.4505 | 48430.0807
| Plasma Stab 900 T30a |  |  | G5060921052.D | Sample |  | 2021-06-10 09:22 | 621.9041 | 8.2251 | 161559.6814 |  | 6.4658 | 48572.1887
| Plasma Stab 900 T30b |  |  | G5060921053.D | Sample |  | 2021-06-10 09:42 | 627.2852 | 8.2252 | 159821.9437 |  | 6.4582 | 47655.6251
| Plasma Stab 900 T30c |  |  | G5060921054.D | Sample |  | 2021-06-10 10:02 | 572.6228 | 8.2251 | 150997.7516 |  | 6.4582 | 49116.9149
| Hep5 CC8 |  |  | G5060921055.D | QC | 8 | 2021-06-10 10:22 | 192.7736 | 8.2252 | 56405.0464 | 96.3868 | 6.4659 | 49807.2156
| Plasma Stab 900 T60a |  |  | G5060921056.D | Sample |  | 2021-06-10 10:42 | 686.2234 | 8.2251 | 176339.1352 |  | 6.4582 | 48246.3209
| Plasma Stab 900 T60b |  |  | G5060921057.D | Sample |  | 2021-06-10 11:01 | 732.4808 | 8.2252 | 195318.65 |  | 6.4583 | 50191.8297
| Plasma Stab 900 T60c |  |  | G5060921058.D | Sample |  | 2021-06-10 11:21 | 763.8077 | 8.219 | 195176.107 |  | 6.4513 | 48172.7676
| Plasma Stab 900 T120a |  |  | G5060921059.D | Sample |  | 2021-06-10 11:41 | 803.3954 | 8.2252 | 203289.1371 |  | 6.4583 | 47788.2378
| Plasma Stab 900 T120b |  |  | G5060921060.D | Sample |  | 2021-06-10 12:01 | 894.5916 | 8.219 | 222711.7514 |  | 6.4513 | 47183.1808
| Plasma Stab 900 T120c |  |  | G5060921061.D | Sample |  | 2021-06-10 12:21 | 827.4442 | 8.2191 | 215368.1682 |  | 6.4505 | 49205.6985
| Plasma Stab 900 T240a |  |  | G5060921062.D | Sample |  | 2021-06-10 12:41 | 918.3097 | 8.219 | 244311.9681 |  | 6.4505 | 50463.1037
| Plasma Stab 900 T240b |  |  | G5060921063.D | Sample |  | 2021-06-10 13:01 | 787.5123 | 8.2252 | 201256.1373 |  | 6.4582 | 48230.8978
| Plasma Stab 900 T240c |  |  | G5060921064.D | Sample |  | 2021-06-10 13:21 | 961.8174 | 8.2191 | 249863.0416 |  | 6.4505 | 49342.905
| Hep5 CC4 |  |  | G5060921065.D | QC | 4 | 2021-06-10 13:41 | 9.4288 | 8.2252 | 10426.0457 | 31.4294 | 6.4582 | 53408.9946

## 273_stab
Sample | [7 empty] | 273 Results |  |  |  | MFOET (ISTD) Results
| Name | Comment | Info. | Data File | Type | Level | Acq. Date-Time | Final Conc. | RT | Area | Accuracy | RT | Resp. |  |  |  |  | 273 HC Stability
| Spike |  |  | G5060921001.D | MatrixBlank |  | 2021-06-09 16:25 | 126.2979 | 9.0721 | 37764.6369 |  | 6.4512 | 48954.0392 | Hep Clearance Time Course Sample Results |  |  |  |  | % of Time 0
| Spike |  |  | G5060921002.D | MatrixBlank |  | 2021-06-09 16:45 | 0 | 9.2062 | 3750.9176 |  | 6.4513 | 48159.3935 | Sample ID | Peak Area | Average | Standard Deviation | Time | Replicate | Average | SD
| Spike |  |  | G5060921003.D | MatrixBlank |  | 2021-06-09 17:05 | 0 | 9.2063 | 2563.0338 |  | 6.4506 | 49705.5305 | T0 rep 1 | 204858.68 | 201908.386 | 6852.8572 | 0 | 101.5% | 100% | 0.03394
| Spike |  |  | G5060921004.D | MatrixBlank |  | 2021-06-09 17:25 | 0 | 9.2609 | 4673.7634 |  | 6.4513 | 46806.429 | T0 rep 2 | 194074.5068 |  |  |  | 96.12%
| Spike |  |  | G5060921005.D | MatrixBlank |  | 2021-06-09 17:45 | 0 | 9.2002 | 1810.3173 |  | 6.4582 | 47636.5873 | T0 rep 3 | 206791.9712 |  |  |  | 102.4%
| Hep5 CC15 |  |  | G5060921006.D | QC | 15 | 2021-06-09 18:04 | 4820.4108 | 9.2062 | 1011536.6647 | 96.4082 | 6.459 | 48065.3218 | T15 rep 1 | 180959.3412 | 171916.9871 | 12129.5182 | 15 | 89.62% | 85.15% | 0.060074
| Hep5 CC15 |  |  | G5060921007.D | QC | 15 | 2021-06-09 18:24 | 4767.927 | 9.2063 | 1019619.2687 | 95.3585 | 6.4582 | 48976.9109 | T15 rep 2 | 158132.8685 |  |  |  | 78.32%
| Hep5 CC15 |  |  | G5060921008.D | QC | 15 | 2021-06-09 18:44 | 4647.3136 | 9.2062 | 1011422.2713 | 92.9463 | 6.4582 | 49830.0319 | T15 rep 3 | 176658.7516 |  |  |  | 87.49%
| Hep5 CC15 |  |  | G5060921009.D | QC | 15 | 2021-06-09 19:04 | 4557.3961 | 9.2063 | 989066.8056 | 91.1479 | 6.4582 | 49679.1447 | T30 rep 1 | 170998.4945 | 165825.7172 | 5108.1236 | 30 | 84.69% | 82.13% | 0.025299
| Hep5 CC15 |  |  | G5060921010.D | Cal | 15 | 2021-06-09 19:24 | 4933.2615 | 9.2062 | 990491.7276 | 98.6652 | 6.4582 | 45999.9838 | T30 rep 2 | 165693.8566 |  |  |  | 82.06%
| Hep5 CC14 |  |  | G5060921011.D | Cal | 14 | 2021-06-09 19:44 | 5119.7663 | 9.2063 | 1049999.7746 | 146.279 | 6.4659 | 47005.2191 | T30 rep 3 | 160784.8004 |  |  |  | 79.63%
| Hep5 CC13 |  |  | G5060921012.D | Cal | 13 | 2021-06-09 20:04 | 2637.3395 | 9.2062 | 568818.3001 | 105.4936 | 6.4667 | 48966.5011 | T60 rep 1 | 188470.4562 | 177570.4239 | 16440.6551 | 60 | 93.34% | 87.95% | 0.081426 | Time (min) | % of Time 0
| Hep5 CC12 |  |  | G5060921013.D | Cal | 12 | 2021-06-09 20:24 | 0 | 9.2063 | 1594.9355 | 0 | 6.4582 | 52771.4594 | T60 rep 2 | 185580.9839 |  |  |  | 91.91% |  |  | 0 | 101.5% | 96.12% | 102.4%
| Hep5 CC11 |  |  | G5060921014.D | Cal | 11 | 2021-06-09 20:44 | 868.8409 | 9.2062 | 194451.3382 | 108.6051 | 6.4582 | 48876.3755 | T60 rep 3 | 158659.8315 |  |  |  | 78.58% |  |  | 15 | 89.62% | 78.32% | 87.49%
| Hep5 CC10 |  |  | G5060921015.D | Cal | 10 | 2021-06-09 21:04 | 424.7072 | 9.2063 | 95366.4117 | 84.9414 | 6.4659 | 46288.9036 | T120 rep 1 | 185708.531 | 143231.3024 | 40388.1423 | 120 | 91.98% | 70.94% | 0.20003 | 30 | 84.69% | 82.06% | 79.63%
| Hep5 CC9 |  |  | G5060921016.D | Cal | 9 | 2021-06-09 21:24 | 288.682 | 9.2062 | 71826.7591 | 82.4806 | 6.459 | 48770.2368 | T120 rep 2 | 138664.9951 |  |  |  | 68.68% |  |  | 60 | 93.34% | 91.91% | 78.58%
| Hep5 CC8 |  |  | G5060921017.D | Cal | 8 | 2021-06-09 21:44 | 202.8866 | 9.2063 | 51813.0047 | 101.4433 | 6.4659 | 47008.1778 | T120 rep 3 | 105320.381 |  |  |  | 52.16% |  |  | 120 | 91.98% | 68.68% | 52.16%
| Hep5 CC7 |  |  | G5060921018.D | Cal | 7 | 2021-06-09 22:04 | 116.4646 | 9.2062 | 35076.0736 | 93.1717 | 6.4582 | 48117.9084 | T240 rep 1 | 122490.7964 | 96442.9003 | 22634.0398 | 240 | 60.67% | 47.77% | 0.1121 | 240 | 60.67% | 40.4% | 42.23%
| Hep5 CC6 |  |  | G5060921019.D | Cal | 6 | 2021-06-09 22:24 | 42.1408 | 9.2063 | 19394.5603 | 52.6761 | 6.4582 | 47540.2486 | T240 rep 2 | 81566.902 |  |  |  | 40.4%
| Hep5 CC5 |  |  | G5060921020.D | Cal | 5 | 2021-06-09 22:44 | 6.877 | 9.2123 | 11590.6718 | 13.7541 | 6.4667 | 45336.5558 | T240 rep 3 | 85271.0025 |  |  |  | 42.23%
| Hep5 CC4 |  |  | G5060921021.D | Cal | 4 | 2021-06-09 23:04 | 0 | 9.2063 | 8903.2885 | 0 | 6.4659 | 46334.7654
| Hep5 CC3 |  |  | G5060921022.D | Cal | 3 | 2021-06-09 23:24 | 0 | 9.2062 | 6064.0567 | 0 | 6.4658 | 47432.6339
| Hep5 CC2 |  |  | G5060921023.D | Cal | 2 | 2021-06-09 23:44 | 190.4135 | 9.2063 | 3068.7076 | 1586.7793 | 6.4505 | 2927.2003
| Hep5 CC1 |  |  | G5060921024.D | Cal | 1 | 2021-06-10 00:03 | 0 | 9.2122 | 5738.1931 | 0 | 6.4513 | 29362.9715
| Spike |  |  | G5060921025.D | MatrixBlank |  | 2021-06-10 00:23 | 0 | 9.2063 | 3325.8199 |  | 6.4582 | 48509.0713
| Hep5 CC9 |  |  | G5060921105.D | QC | 9 | 2021-06-11 02:58 | 294.0059 | 9.2063 | 79764.1973 | 84.0017 | 6.4582 | 53327.1543
| Plasma Stab 273 T0a |  |  | G5060921106.D | Sample |  | 2021-06-11 03:18 | 925.0577 | 9.2062 | 204858.68 |  | 6.459 | 48530.5768
| Plasma Stab 273 T0b |  |  | G5060921107.D | Sample |  | 2021-06-11 03:38 | 871.4871 | 9.2063 | 194074.5068 |  | 6.4582 | 48641.9217
| Plasma Stab 273 T0c |  |  | G5060921108.D | Sample |  | 2021-06-11 03:58 | 916.5508 | 9.2062 | 206791.9712 |  | 6.4513 | 49418.6969
| Plasma Stab 273 T15a |  |  | G5060921109.D | Sample |  | 2021-06-11 04:18 | 806.3203 | 9.2063 | 180959.3412 |  | 6.4582 | 48797.029
| Plasma Stab 273 T15b |  |  | G5060921110.D | Sample |  | 2021-06-11 04:38 | 717.7085 | 9.2062 | 158132.8685 |  | 6.4582 | 47548.755
| Plasma Stab 273 T15c |  |  | G5060921111.D | Sample |  | 2021-06-11 04:58 | 764.1866 | 9.2063 | 176658.7516 |  | 6.4582 | 50095.5516
| Plasma Stab 273 T30a |  |  | G5060921112.D | Sample |  | 2021-06-11 05:18 | 766.4824 | 9.2062 | 170998.4945 |  | 6.4582 | 48354.4996
| Plasma Stab 273 T30b |  |  | G5060921113.D | Sample |  | 2021-06-11 05:38 | 714.3279 | 9.2063 | 165693.8566 |  | 6.4582 | 50041.9505
| Plasma Stab 273 T30c |  |  | G5060921114.D | Sample |  | 2021-06-11 05:58 | 711.8305 | 9.2062 | 160784.8004 |  | 6.459 | 48718.0444
| Hep5 CC12 |  |  | G5060921115.D | QC | 12 | 2021-06-11 06:18 | 0 | 9.2002 | 6087.047 | 0 | 6.4582 | 64345.1052
| Plasma Stab 273 T60a |  |  | G5060921116.D | Sample |  | 2021-06-11 06:38 | 788.0218 | 9.2062 | 188470.4562 |  | 6.4582 | 51929.1219
| Plasma Stab 273 T60b |  |  | G5060921117.D | Sample |  | 2021-06-11 06:57 | 778.3748 | 9.2063 | 185580.9839 |  | 6.4582 | 51726.8027
| Plasma Stab 273 T60c |  |  | G5060921118.D | Sample |  | 2021-06-11 07:17 | 648.8324 | 9.2062 | 158659.8315 |  | 6.459 | 52393.6294
| Plasma Stab 273 T120a |  |  | G5060921119.D | Sample |  | 2021-06-11 07:37 | 773.1904 | 9.2063 | 185708.531 |  | 6.4582 | 52087.4395
| Plasma Stab 273 T120b |  |  | G5060921120.D | Sample |  | 2021-06-11 07:57 | 589.8723 | 9.2062 | 138664.9951 |  | 6.459 | 49994.9129
| Plasma Stab 273 T120c |  |  | G5060921121.D | Sample |  | 2021-06-11 08:17 | 441.9397 | 9.2063 | 105320.381 |  | 6.4582 | 49338.025
| Plasma Stab 273 T240a |  |  | G5060921122.D | Sample |  | 2021-06-11 08:37 | 509.3703 | 9.2062 | 122490.7964 |  | 6.459 | 50492.9599
| Plasma Stab 273 T240b |  |  | G5060921123.D | Sample |  | 2021-06-11 08:57 | 320.2105 | 9.2063 | 81566.902 |  | 6.4582 | 50696.4302
| Plasma Stab 273 T240c |  |  | G5060921124.D | Sample |  | 2021-06-11 09:17 | 327.3719 | 9.2062 | 85271.0025 |  | 6.459 | 51999.0199
| Hep5 CC7 |  |  | G5060921125.D | QC | 7 | 2021-06-11 09:37 | 119.3182 | 9.2063 | 40320.5146 | 95.4546 | 6.4582 | 54392.6895

## HC_G5_Ametryn_MDL
MDL Calculation
Ametryn
Sample | [6 empty] | Ametryn Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Final Conc. | Accuracy | Area
| HC_G5 CC 5 |  | 7010_1_071421060.D | Cal | 5 | 2021-07-15 14:41 | 16.8328 | 49.8674 | 99.7349 | 15990.395
| HC_G5 CC 5 |  | 7010_1_071421061.D | Cal | 5 | 2021-07-15 15:04 | 16.8372 | 53.2846 | 106.5691 | 16875.5731
| HC_G5 CC 5 |  | 7010_1_071421062.D | Cal | 5 | 2021-07-15 15:28 | 16.8328 | 50.4385 | 100.877 | 16342.8132
| HC_G5 CC 5 |  | 7010_1_071421063.D | Cal | 5 | 2021-07-15 15:52 | 16.8372 | 49.7643 | 99.5287 | 16116.5717
| HC_G5 CC 5 |  | 7010_1_071421064.D | Cal | 5 | 2021-07-15 16:16 | 16.8372 | 47.5132 | 95.0265 | 15251.7184
| HC_G5 CC 5 |  | 7010_1_071421065.D | Cal | 5 | 2021-07-15 16:40 | 16.8415 | 49.7923 | 99.5846 | 15816.386
| HC_G5 CC 5 |  | 7010_1_071421066.D | Cal | 5 | 2021-07-15 17:04 | 16.8287 | 49.3396 | 98.6792 | 15306.5173
[7 empty] | sd | 1.718
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  | MDL (nM) | 5.4
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [6 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf
Sample | [8 empty] | Ametryn Method | Ametryn Results | [6 empty] | Qualifier (212.0 -> 71.1) Results |  | Qualifier (212.0 -> 122.0) Results |  | 4NT13C6 (ISTD) Results |  | Qualifier (143.0 -> 84.1) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | Ratio | MI | RT | Resp. | Ratio | MI
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421060.D | Cal | 5 | 2021-07-15 14:41 | 50 | 16.8328 | 15990.395 | FALSE | 49.8674 |  | 49.8674 | 99.7349 | 67.64 | FALSE | 83.1989 | FALSE | 11.2831 | 1370707.8045 | 57.8704 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421061.D | Cal | 5 | 2021-07-15 15:04 | 50 | 16.8372 | 16875.5731 | FALSE | 53.2846 |  | 53.2846 | 106.5691 | 63.9641 | FALSE | 73.065 | FALSE | 11.2831 | 1353816.8951 | 57.6503 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421062.D | Cal | 5 | 2021-07-15 15:28 | 50 | 16.8328 | 16342.8132 | FALSE | 50.4385 |  | 50.4385 | 100.877 | 65.8591 | FALSE | 76.3819 | FALSE | 11.2831 | 1385055.6785 | 58.0017 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421063.D | Cal | 5 | 2021-07-15 15:52 | 50 | 16.8372 | 16116.5717 | FALSE | 49.7643 |  | 49.7643 | 99.5287 | 65.5413 | FALSE | 77.7659 | FALSE | 11.2831 | 1384386.1777 | 58.7338 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421064.D | Cal | 5 | 2021-07-15 16:16 | 50 | 16.8372 | 15251.7184 | FALSE | 47.5132 |  | 47.5132 | 95.0265 | 74.3781 | FALSE | 83.601 | FALSE | 11.2831 | 1372166.9626 | 57.4697 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421065.D | Cal | 5 | 2021-07-15 16:40 | 50 | 16.8415 | 15816.386 | FALSE | 49.7923 |  | 49.7923 | 99.5846 | 68.3005 | FALSE | 86.2636 | FALSE | 11.2831 | 1357836.955 | 57.5271 | FALSE
| HC_G5 CC 5 | E:\Shark Tank\RawData\Ametryn071421 |  | 62 | 7010_1_071421066.D | Cal | 5 | 2021-07-15 17:04 | 50 | 16.8287 | 15306.5173 | FALSE | 49.3396 |  | 49.3396 | 98.6792 | 74.2196 | FALSE | 78.0108 | FALSE | 11.2833 | 1326121.6178 | 57.4054 | FALSE

## MDL Data
Sample | [8 empty] | 273 Method | 273 Results | [6 empty] | Qualifier (100.0 -> 69.0) Results |  | Qualifier (119.0 -> 69.0) Results |  | MFOET (ISTD) Results |  | Qualifier (448.0 -> 96.0) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | Ratio | MI | RT | Resp. | Ratio | MI
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021131.D | Cal | 5 | 2021-07-02 13:49 | 50 | 9.2001 | 15633.4976 | TRUE | 53.5563 |  | 53.5563 | 107.1126 | 37.3363 | FALSE | 49.1411 | FALSE | 6.4513 | 24147.0443 | 92.2264 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021132.D | Cal | 5 | 2021-07-02 14:09 | 50 | 9.2002 | 14791.6408 | TRUE | 50.1808 |  | 50.1808 | 100.3616 | 44.522 | FALSE | 55.4969 | FALSE | 6.4505 | 24383.5674 | 82.4655 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021133.D | Cal | 5 | 2021-07-02 14:29 | 50 | 9.2001 | 15179.3903 | TRUE | 50.5731 |  | 50.5731 | 101.1461 | 36.4328 | FALSE | 54.3938 | FALSE | 6.4513 | 24828.6705 | 79.3475 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021134.D | Cal | 5 | 2021-07-02 14:49 | 50 | 9.2002 | 13469.5553 | TRUE | 43.6495 |  | 43.6495 | 87.2991 | 46.4946 | FALSE | 64.9793 | FALSE | 6.4505 | 25526.5462 | 82.4326 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021135.D | Cal | 5 | 2021-07-02 15:09 | 50 | 9.2001 | 15240.2234 | TRUE | 52.1521 |  | 52.1521 | 104.3042 | 37.2569 | FALSE | 59.2005 | FALSE | 6.4513 | 24173.4202 | 85.086 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021136.D | Cal | 5 | 2021-07-02 15:29 | 50 | 9.2002 | 15818.9618 | TRUE | 52.7564 |  | 52.7564 | 105.5128 | 33.5243 | FALSE | 61.5233 | FALSE | 6.4505 | 24803.9699 | 80.159 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021137.D | Cal | 5 | 2021-07-02 15:49 | 50 | 9.2001 | 14111.2292 | TRUE | 47.1318 |  | 47.1318 | 94.2636 | 39.5276 | FALSE | 62.582 | FALSE | 6.4513 | 24766.7712 | 76.279 | FALSE
Sample | [8 empty] | 900 Method | 900 Results | [6 empty] | Qualifier (182.0 -> 131.1) Results |  | MFOET (ISTD) Results |  | Qualifier (448.0 -> 96.0) Results
| Name | Data Path | Comment | Vial | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | MI | Calc. Conc. | Custom Calc. | Final Conc. | Accuracy | Ratio | MI | RT | Resp. | Ratio | MI
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021131.D | Cal | 5 | 2021-07-02 13:49 | 50 | 8.2129 | 15007.7643 | FALSE | 52.8096 |  | 52.8096 | 105.6191 | 38.1647 | FALSE | 6.4513 | 24147.0443 | 92.2264 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021132.D | Cal | 5 | 2021-07-02 14:09 | 50 | 8.2191 | 15223.8438 | FALSE | 53.0503 |  | 53.0503 | 106.1005 | 34.8798 | FALSE | 6.4505 | 24383.5674 | 82.4655 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021133.D | Cal | 5 | 2021-07-02 14:29 | 50 | 8.219 | 14734.903 | FALSE | 50.426 |  | 50.426 | 100.8519 | 38.4297 | FALSE | 6.4513 | 24828.6705 | 79.3475 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021134.D | Cal | 5 | 2021-07-02 14:49 | 50 | 8.2191 | 15185.8315 | FALSE | 50.5484 |  | 50.5484 | 101.0967 | 36.8502 | FALSE | 6.4505 | 25526.5462 | 82.4326 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021135.D | Cal | 5 | 2021-07-02 15:09 | 50 | 8.219 | 14592.0505 | FALSE | 51.2907 |  | 51.2907 | 102.5814 | 39.178 | FALSE | 6.4513 | 24173.4202 | 85.086 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021136.D | Cal | 5 | 2021-07-02 15:29 | 50 | 8.2191 | 12955.9401 | FALSE | 44.3821 |  | 44.3821 | 88.7643 | 40.6522 | FALSE | 6.4505 | 24803.9699 | 80.159 | FALSE
| Hep5 CC5 | E:\Shark Tank\RawData\900 |  | 12 | G5063021137.D | Cal | 5 | 2021-07-02 15:49 | 50 | 8.219 | 13843.2661 | FALSE | 47.493 |  | 47.493 | 94.986 | 42.4494 | FALSE | 6.4513 | 24766.7712 | 76.279 | FALSE

## 900 MDL
MDL Calculation
900
Sample | [6 empty] | 900 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Final Conc. | Accuracy | Area
| Hep5 CC5 |  | G5063021131.D | Cal | 5 | 2021-07-02 13:49 | 8.2129 | 52.8096 | 105.6191 | 15007.7643
| Hep5 CC5 |  | G5063021132.D | Cal | 5 | 2021-07-02 14:09 | 8.2191 | 53.0503 | 106.1005 | 15223.8438
| Hep5 CC5 |  | G5063021133.D | Cal | 5 | 2021-07-02 14:29 | 8.219 | 50.426 | 100.8519 | 14734.903
| Hep5 CC5 |  | G5063021134.D | Cal | 5 | 2021-07-02 14:49 | 8.2191 | 50.5484 | 101.0967 | 15185.8315
| Hep5 CC5 |  | G5063021135.D | Cal | 5 | 2021-07-02 15:09 | 8.219 | 51.2907 | 102.5814 | 14592.0505
| Hep5 CC5 |  | G5063021136.D | Cal | 5 | 2021-07-02 15:29 | 8.2191 | 44.3821 | 88.7643 | 12955.9401
| Hep5 CC5 |  | G5063021137.D | Cal | 5 | 2021-07-02 15:49 | 8.219 | 47.493 | 94.986 | 13843.2661
[7 empty] | sd | 3.087
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  | MDL (nM) | 9.7
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [6 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 273 MDL
MDL Calculation
273
Sample | [6 empty] | 273 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Final Conc. | Accuracy | Area
| Hep5 CC5 |  | G5063021131.D | Cal | 5 | 2021-07-02 13:49 | 9.2001 | 53.5563 | 107.1126 | 15633.4976
| Hep5 CC5 |  | G5063021132.D | Cal | 5 | 2021-07-02 14:09 | 9.2002 | 50.1808 | 100.3616 | 14791.6408
| Hep5 CC5 |  | G5063021133.D | Cal | 5 | 2021-07-02 14:29 | 9.2001 | 50.5731 | 101.1461 | 15179.3903
| Hep5 CC5 |  | G5063021134.D | Cal | 5 | 2021-07-02 14:49 | 9.2002 | 43.6495 | 87.2991 | 13469.5553
| Hep5 CC5 |  | G5063021135.D | Cal | 5 | 2021-07-02 15:09 | 9.2001 | 52.1521 | 104.3042 | 15240.2234
| Hep5 CC5 |  | G5063021136.D | Cal | 5 | 2021-07-02 15:29 | 9.2002 | 52.7564 | 105.5128 | 15818.9618
| Hep5 CC5 |  | G5063021137.D | Cal | 5 | 2021-07-02 15:49 | 9.2001 | 47.1318 | 94.2636 | 14111.2292
[7 empty] | sd | 3.507
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  | MDL (nM) | 11.02
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [6 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## ValueList_Helper (hidden)
Sample
Blank
Cal
QC
CC
DoubleBlank
MatrixSpike
MatrixSpikeDup
MatrixBlank
TuneCheck
ResponseCheck